SERS Raman Data Analysis Report

R6G Sensor Development — Silver Nanoparticle SERS Substrate

Data acquired 2026-07-04  |  B&W Tek BWS465-785H Raman spectrometer (785 nm)  |  Report generated 2026-08-10

# 1. Overview

The selected folder contains 78 Raman spectroscopy CSV files (16 MB total) collected in a single session on 2026-07-04 with a B&W Tek BWS465-785H portable Raman spectrometer (785 nm excitation, 2048-pixel detector, Raman shift range − 46 to 2842 cm⁻¹). The files document development of a surface-enhanced Raman scattering (SERS) sensor: silver nanoparticles are synthesized on a substrate via nicotinamide (NAM) reduction of AgNO3, and rhodamine 6G (R6G) is used as the test analyte to characterize sensitivity across a six-decade concentration range (1 nM to 1 mM).

Every file shares the same BWRam4.15_8 export format: an instrument metadata header (~70 fields) followed by a per-pixel table (Pixel, Wavelength, Wavenumber, Raman Shift, Dark, Reference, Raw data, Dark Subtracted, plus derived columns). All analysis below uses the Dark Subtracted trace.

Figure 1. Five blank Si-substrate spectra overlaid, showing the reproducible 520 cm⁻¹ silicon peak and flat baseline.

# 2. Folder-by-Folder Summary

| Folder | Files | Description | Integration |
| --- | --- | --- | --- |
| substarte/ | 5 | Blank Si substrate, no analyte, no NP coating | 10 s x4 acc |
| namon5hf6mgagno3/ | 7 | Ag nanoparticle synthesis test: nicotinamide (NAM) reduction of AgNO3 at 0.1 and 1 mg/mL NAM | 10 s x4 acc |
| r6g/ | 6 | Early R6G-on-substrate trials (6 single spots) | 10 s x4 acc |
| r6g-sensor/ | 17 | R6G concentration series: 1 nM – 1 mM, replicate spectra per level | 10 s x4 acc |
| rawr6g/ | 39 | Sequential single-spot acquisitions on one R6G sample while sweeping accumulation count (1–15) and integration time | 1–15 s, 1–15 acc |
| SP_13–SP_16.csv (root) | 4 | Earliest measurements of the session (10:58–11:09 AM), likely preliminary/setup scans | 20 s x2 acc |

# 3. Key Findings

## 3.1 Blank substrate is clean and reproducible

The five blank substrate scans overlay almost perfectly (see Figure 1 above), with a flat near-zero baseline and a single sharp peak at ~520 cm⁻¹. This is the characteristic first-order phonon mode of crystalline silicon, confirming the measurements were taken on a Si wafer substrate. Low run-to-run variability in the blanks indicates good instrument stability and a suitable reference for background subtraction.

## 3.2 R6G is clearly detected down to 1 nM, with a non-monotonic concentration response

Average spectra for each of the six R6G concentration levels (1 nM – 1 mM) show the diagnostic R6G SERS bands — in-plane ring modes near 612 and 774 cm⁻¹, C–C stretching modes at 1310, 1360, and 1509 cm⁻¹, and the ring-stretch mode near 1650 cm⁻¹ — superimposed on the silicon 520 cm⁻¹ substrate line.

Figure 2. Average R6G spectra by concentration, fingerprint region (traces offset for clarity). Dotted lines mark reference R6G SERS peak positions.

Baseline-corrected peak heights (table below, Figure 3) increase monotonically from 1 nM through 10 µM, then drop at 1 mM. This is a well-documented SERS behavior: at high analyte concentration, molecules can form multilayers or aggregates on the enhancing surface, so molecules beyond the first layer sit outside the ~1–10 nm electromagnetic “hot-zone” and no longer contribute proportionally to signal, while self-quenching/reabsorption can further reduce yield. Practically, this indicates 10 µM is close to the substrate's optimal working concentration for this NP batch, and the usable dynamic range for a monotonic calibration curve is roughly 1 nM – 10 µM (4 orders of magnitude).

Figure 3. Baseline-corrected intensity of eight R6G marker peaks vs. concentration (log x-axis).

## 3.2.1 Peak intensity data (baseline-corrected counts)

| Sample | Conc. (M) | n | 612 | 774 | 1183 | 1310 | 1360 | 1509 | 1575 | 1650 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| r6g-1nM | 1e-9 | 2 | 15 | 14 | 19 | 12 | 27 | 31 | 17 | 16 |
| r6g-10nM | 1e-8 | 3 | 27 | 5 | 21 | 21 | 10 | 24 | 20 | 25 |
| r6g-100nM | 1e-7 | 3 | 59 | 28 | 37 | 53 | 76 | 84 | 31 | 32 |
| r6g-1microM | 1e-6 | 5 | 138 | 74 | 73 | 98 | 200 | 217 | 18 | 63 |
| r6g-10microM | 1e-5 | 2 | 233 | 126 | 103 | 109 | 333 | 345 | 30 | 93 |
| r6g-1mM | 1e-3 | 2 | 113 | 57 | 70 | 44 | 166 | 181 | 14 | 57 |

## 3.3 Lower NAM concentration gives a stronger nanoparticle SERS signature

Comparing the two NAM/AgNO3 substrate-synthesis conditions, the 0.1 mg/mL NAM batch shows a distinctly stronger, broader feature around 230–250 cm⁻¹ than the 1 mg/mL batch (Figure 4). This low-wavenumber region is where Ag–ligand / Ag–N stretching modes typically appear for molecules bound to freshly formed silver nanoparticle surfaces, so its greater intensity at the lower NAM concentration suggests more/better-exposed nanoparticle surface sites — i.e., 0.1 mg/mL NAM may be the more effective reducing-agent concentration for this substrate recipe. Both conditions still show the underlying 520 cm⁻¹ Si peak, confirming both were measured on the same substrate type.

Figure 4. NAM/AgNO3 substrate spectra, 0.1 mg/mL (blue) vs 1 mg/mL (red).

## 3.4 rawr6g/ is a parameter-sweep, not a clean replicate set

The 39 rawr6g files were captured back-to-back over about 15 minutes (20:02–20:17) on what appears to be a single R6G spot, while the operator swept accumulation count (“average number”) from 1 up to 15 and integration time from 1 to 15 s. Max peak intensity jumps in discrete steps rather than drifting smoothly (Figure 5), which lines up with parameter changes rather than sample or laser instability, though a genuine spot-to-spot or photobleaching effect cannot be ruled out without the full parameter log. This set is best read as an acquisition-settings optimization run rather than a concentration or time-series dataset, and its intensities should not be directly compared to r6g-sensor/ without normalizing for accumulation count and integration time.

Figure 5. Max peak intensity across the 39 sequential rawr6g acquisitions, showing step changes consistent with parameter sweeps.

# 4. Recommendations

- Use the r6g-sensor/ series (1 nM–10 µM) as the calibration set for any LOD/sensitivity claims — it is the only concentration series with a monotonic, physically interpretable trend.

- Flag 1 mM as outside the linear/monotonic range; if higher-concentration data is needed, consider diluting or reporting it separately with an aggregation/multilayer caveat.

- For substrate fabrication, prioritize repeating the 0.1 mg/mL NAM condition with more replicates — it shows the stronger Ag-related SERS feature versus 1 mg/mL, but currently only 3 vs 4 files support the comparison.

- Re-run or clearly relabel rawr6g/ with the accumulation/integration parameter for each file so it can be used quantitatively; as named, file order is the only clue to acquisition settings.

- Increase replicate counts for the thinner groups (1 nM n=2, 10 µM n=2, 1 mM n=2) to strengthen the calibration curve's statistical confidence.

# 5. Files Produced

Supporting analysis outputs (peak intensity table, summary JSON, and all figures) accompany this report. Raw parsing used the Dark Subtracted column from each CSV; peak heights are local maxima within ±8 cm⁻¹ of literature R6G band positions, minus a local median baseline from ±(40–48) cm⁻¹ flanking windows.