# Boronic Acid as the Capture Chemistry — Design Rationale

Project: SERS nanobiosensor for rapid bacterial identification in sepsis Prepared: August 2026

## 1. Summary

Boronic acid is proposed as the bacterial capture element on the Ag/Si SERS substrate. It forms reversible covalent bonds with cis-diols, which are abundant on bacterial cell-wall glycans, and it can be immobilised on silver as a self-assembled monolayer that doubles as a Raman reporter.

The approach is well precedented. Reported detection limits using 4-mercaptophenylboronic acid (4-MPBA) reach 8–13 CFU mL⁻¹ for respiratory pathogens on a magnetic SERS strip [1], and capture efficiencies of 99.65% at 10³ CFU mL⁻¹ on 4-MPBA-functionalised silver dendrites [2].

Two limitations must be designed around: competing cis-diols in blood (principally glucose) and the pH mismatch between simple phenylboronic acids and physiological conditions. Both have established solutions.

## 2. The chemistry

Boronic acids condense with 1,2- or 1,3-cis-diols to form cyclic boronate esters, releasing water. The reaction is reversible and pH-dependent: binding is stronger above the pKa of the boronic acid, where the tetrahedral boronate anion predominates over the neutral trigonal form [3].

R–B(OH)₂  +  HO–CH–CH–OH   ⇌   cyclic boronate ester  +  2 H₂O

This gives three properties that matter for a diagnostic device:

- synthetic recognition — no antibodies, no cold chain, no batch variability

- reversibility — the surface can be regenerated by lowering pH

- broad glycan capture — one ligand captures many bacterial species

## 3. Why 4-mercaptophenylboronic acid specifically

4-MPBA is bifunctional in exactly the way this device needs. The thiol chemisorbs onto silver or gold, forming an ordered monolayer with the boronic acid presented outward for capture [1].

More usefully, 4-MPBA is itself a strong Raman reporter. Its SERS spectrum shows characteristic bands at 1072 cm⁻¹ (C–S-stretch-coupled ring breathing) and 1585 cm⁻¹ (aromatic C=C) [4]. Binding events perturb these bands, so the same molecule performs capture and signal transduction.

This matters practically: it gives an internal reference. A change in the ratio of reporter bands to analyte bands indicates binding, independent of absolute intensity — which is the usual weakness of SERS quantification, since hotspot density varies between substrates and between spots.

## 4. What boronic acid actually binds on bacteria — an important nuance

The literature commonly states that boronic acid captures bacteria by binding peptidoglycan [1]. More careful studies attribute the interaction primarily to lipopolysaccharide in Gram-negative organisms and teichoic acids in Gram-positive organisms [5, 6].

The structural argument favours the second interpretation, and it is directly relevant to our NAM work.

Boronate ester formation requires two hydroxyls in a suitable cis geometry. In the peptidoglycan glycan backbone:

| Position | N-acetylmuramic acid | N-acetylglucosamine |
| --- | --- | --- |
| C-1 | glycosidic linkage (in polymer) | glycosidic linkage |
| C-2 | acetamido — not a hydroxyl | acetamido — not a hydroxyl |
| C-3 | D-lactyl ether — blocked | free OH |
| C-4 | glycosidic linkage | free OH |

In NAM, C-2 carries the acetamido group and C-3 the lactyl ether, so no 1,2-diol is available. In GlcNAc, the C-3 and C-4 hydroxyls exist but are trans-diequatorial, a geometry that binds boronic acid weakly. In the intact polymer, C-1 and C-4 are further consumed by the β-1,4 glycosidic linkages.

Conclusion: the peptidoglycan glycan backbone is a poor boronic acid target. Capture on intact bacteria is far more likely occurring through LPS — whose O-antigen sugars and Kdo/heptose inner core are rich in cis-diols — and through wall and lipoteichoic acids on Gram-positives. Recent work reports a 40-fold binding enhancement from bivalent boron conjugates acting mainly on wall teichoic acid [5].

### Why this is worth stating explicitly

It means the capture chemistry and the NAM detection chemistry target different parts of the cell. That is a deliberate design choice, not an oversight, and framing it that way is better than letting a reviewer notice it.

It also makes a testable prediction: capture efficiency should differ systematically between Gram-positive and Gram-negative organisms, because the diol density and accessibility differ. Boronic acid-modified dendrimers have already been used to discriminate Gram types on exactly this basis [6] — which both supports the approach and means our contribution must be differentiated from it.

## 5. Limitation 1 — competing cis-diols in blood

Boronic acid binds any cis-diol. Blood contains glucose at roughly 5 mM, plus glycoproteins, ribose-containing nucleotides and catechols.

Reported phenylboronic acid affinities follow the order fructose > arabinose, ribose > galactose > glucose, spanning only one to two orders of magnitude [7]. That selectivity is inadequate when the interferent exceeds the target by many orders of magnitude in concentration.

Mitigations, in order of practicality:

- Sample preparation. Separate or lyse bacteria and wash before capture. This is the most reliable route and is compatible with the microfluidic front-end already planned.

- Boronic acids with poor glucose affinity. Glucose binds relatively weakly among common sugars; ligand choice can widen the margin.

- Multivalent presentation. A bacterium presents many diols simultaneously, so a multivalent surface gains avidity that a monovalent sugar like glucose cannot match. This is the strongest conceptual argument for why the approach works at all in biological media.

- Ratiometric readout. Using the 4-MPBA reporter bands as an internal reference reduces sensitivity to non-specific binding.

## 6. Limitation 2 — pH

Simple phenylboronic acids have pKa near 8.8 [8]. Blood is pH 7.4, below the pKa, where the weakly binding neutral trigonal form dominates. Binding is therefore poor at exactly the pH the device must operate at.

Solutions:

Benzoboroxoles. Ring strain from the fused five-membered ring lowers the pKa to approximately 7.2–7.3 [8, 9]. They bind efficiently at physiological pH and show higher saccharide affinity than other Wulff-type boronic acids. They also bind non-reducing sugars — methyl α-D-mannopyranoside, methyl β-D-galactopyranoside — that simple boronic acids cannot [8].

Fluorinated boronic acids. Electron-withdrawing substituents such as ortho- fluorine lower the pKa. 3-fluoro-4-mercaptophenylboronic acid retains the thiol anchor while shifting the working pH downward.

Wulff-type boronic acids. An ortho-aminomethyl group provides intramolecular B–N interaction, stabilising the tetrahedral form at lower pH.

Recommendation: evaluate a benzoboroxole derivative alongside 4-MPBA. 4-MPBA is the established SERS reporter and the sensible starting point; a benzoboroxole variant is the likely route to working in undiluted physiological fluid.

## 7. How this fits the device architecture

The design separates capture from identification:

| Stage | Element | Question answered |
| --- | --- | --- |
| Capture | Boronic acid monolayer on Ag/Si | Is bacterial material present, and concentrate it at the hotspots |
| Identification | SERS fingerprint + DFT-assigned bands | Gram-positive or Gram-negative? |

This is deliberate. Requiring a single element to be both broadly capturing and highly discriminating is the usual failure mode of affinity biosensors. Here the boronic acid needs only to be broad and robust, while discrimination comes from the vibrational fingerprint — which is where the NAM reference spectrum and the band assignments already completed become directly useful.

Concentrating analyte at the plasmonic surface also addresses the central practical problem in SERS: enhancement falls off steeply with distance from the metal, so a capture layer that holds the target within a few nanometres is worth orders of magnitude in signal.

## 8. Proposed experimental sequence

- Establish reproducible Ag/Si substrates — prerequisite; currently blocked by the contamination issue under investigation.

- Form and verify the 4-MPBA monolayer. Confirm the 1072 and 1585 cm⁻¹ bands are present, stable, and reproducible across substrates. This doubles as a substrate quality-control metric, which the current fabrication work needs anyway.

- Model diol binding. Test with a simple sugar to confirm the reporter bands respond to boronate ester formation.

- Bacterial capture in buffer. Establish capture efficiency and detection limit for one Gram-positive and one Gram-negative organism.

- Compare Gram types. Test the prediction that capture efficiency differs.

- Move to serum, then blood. Quantify the glucose interference and evaluate whether sample preparation or a lower-pKa ligand is required.

- Regeneration. Establish how many bind–release cycles the surface tolerates.

Steps 2 and 3 can begin as soon as substrate reproducibility is restored, and would give an independent handle on substrate quality.

## 9. Open questions

- Does 4-MPBA monolayer formation interfere with the Ag/Si hotspot structure, or passivate the surface?

- What is the actual capture efficiency in whole blood versus buffer?

- Does bound bacterial material sit close enough to the surface for its own Raman bands to be enhanced, or do we only observe the reporter?

- Is boronic acid alone sufficient for the sensitivity required in sepsis, where bacterial loads can be 1–10 CFU mL⁻¹?

- How does the boronic acid layer affect the photostability limits established for NAM powder? Plasmonic substrates concentrate the field and may lower the damage threshold considerably.

## References

- Magnetic SERS strip based on 4-mercaptophenylboronic acid-modified Fe₃O₄@Au for active capture and simultaneous detection of respiratory bacteria. Biosensors 13 (2023) 210.

- Rapid concentration detection and differentiation of bacteria in skimmed milk using SERS mapping on 4-mercaptophenylboronic acid functionalised silver dendrites. Analytical and Bioanalytical Chemistry (2016).

- Structure–reactivity relationships in boronic acid–diol complexation. ACS Omega 4 (2019).

- SERS characterisation of 4-mercaptophenylboronic acid: ring-breathing and aromatic C=C reporter bands.

- An efficient method for site-selective boronation of peptides and proteins applied in magnified bacterial imaging. Nature Communications (2025).

- Rapid and selective discrimination of Gram-positive and Gram-negative bacteria by boronic acid-modified poly(amidoamine) dendrimer. Analytical Chemistry 91 (2019).

- Biosensors with boronic acid-based materials as the recognition elements and signal labels. Sensors (2023).

- Multivalent benzoboroxole functionalised polymers as gp120 glycan-targeted microbicide entry inhibitors. Biomacromolecules (2010).

- A benzoboroxole-based affinity ligand for glycoprotein purification at physiological pH. Journal of Molecular Recognition 29 (2016).

Full citation details should be completed from the source articles before this text is used in a report or manuscript.