# Current Lab Resources & Materials Inventory

Instrument, substrate, chemical, and biomarker stock — as evidenced in project files, for use in deciding next steps

Project: SERS nanobiosensor for rapid bacterial identification in sepsis

Prepared: 14 August 2026

This is a single reference sheet for what is actually in hand right now — instrument, substrate, and chemicals — separated clearly from what is planned but not yet available. It is compiled from the project's own prior audits (Comprehensive Research Progress Report, Document A/B, fabrication and measurement logs), not re-derived from scratch, and every figure below traces to one of those files. Where a spec was never confirmed, that is stated explicitly rather than assumed — those gaps are exactly the kind of thing that should be resolved before committing to a specific next experiment.

# 1. Raman / SERS Instrument (the laser)

| Parameter | Value |
| --- | --- |
| Model | B&W Tek i-Raman Plus — spectrometer unit BWS465-785H, system model BTC665N-785H-SYS |
| Probe / microscope head | BAC102-785E |
| Excitation wavelength | 785 nm (784.92 nm as reported by the instrument's own calibration) |
| Laser class / max power | Class 3B; nominal maximum output 495 mW |
| Objective | 20× Plan objective (used in all measurements reviewed); 50× LMPlan objective available but unused |
| Detector | 2048-pixel array |
| Spectral range | ≈ −46 to 2842 cm⁻¹ |
| Spectral sampling | ≈1.4–1.8 cm⁻¹/pixel, varies by region; mean resolution ≈1.41 cm⁻¹ (one QC dataset) |
| Software | BWSpec / BWRam (v4.15_8 seen in raw file metadata) |
| Settings actually used | Laser power 5–90% of max (unitless "laser_powerlevel" field); integration 3–60 s; 1–30 accumulations, depending on experiment |

Known gap: true laser power at the sample has never been measured in mW with a power meter. The instrument only logs a 0–100 "laser_powerlevel" setting, and different project documents inconsistently call this "percent" or "Watts." One planning note found a general spec sheet figure (~30 mW nominal at probe / ~42 mW at laser port for the BWS465-785x class) but this was not confirmed against this specific unit. This matters for any cross-experiment or cross-paper comparison and is a five-minute fix with a power meter if one is accessible.

# 2. Silicon Substrate

| Parameter | Value |
| --- | --- |
| Crystal type | Single-crystal, Czochralski-grown |
| Doping | n-type, phosphorus-doped |
| Crystal orientation | ⟨100⟩ |
| Resistivity | 1–10 Ω·cm |
| Diameter | 2-inch |
| Thickness | 280 ± 25 µm |
| Polish | Single-side polished |
| Supplier | Not documented in any reviewed file |

This spec was recorded in the project's own "Raman Spectroscopy System and Silicon Wafer Specifications" note and matches the doping type/resistivity band used in the closest literature precedent for the fabrication route in use (Bai et al. 2019, Nanomaterials — MACE + photonic Ag reduction on n-type 1–10 Ω·cm Si, EF ≈1.4×10⁸ with R6G). That match is favourable — it means the published protocol transfers with no wafer substitution needed — but the wafer supplier itself, and independent confirmation of the exact resistivity/orientation on the physical stock in hand, were not found on file.

# 3. Fabrication Chemicals — Ag/Si SERS Substrate (in active use)

Route: one-step metal-assisted chemical etching (MACE), electroless silver deposition. Adapted from the host lab's own prior published protocol (Soni, Singh, Singh & Agarwal, IEEE APSCON 2023).

| Reagent | Role | Concentration / amount actually used | Notes |
| --- | --- | --- | --- |
| Silver nitrate (AgNO₃) | Ag⁺ source for electroless deposition | 6 mg per 45 mL DI water (baseline ≈20.0 mM); DOE series tests 5/10/20/40/80 mM | Supplier/purity not documented |
| Hydrofluoric acid (HF) | Si etchant, combined with AgNO₃ | 5 mL or 10 mL (crossed with 3/4/5/6 min soak) in the 8-condition fabrication matrix | Stock assumed 49% w/w, density 1.15 g/mL — NOT independently confirmed against the reagent bottle |
| Piranha solution | Pre-clean, removes organic contamination | Standard prep | Supplier/purity not documented |
| Buffered HF | Native-oxide strip immediately before deposition | ~5 minute dip | — |
| Deionised / ultrapure water | Primary solvent and rinse | — | Selected over PBS/DMSO/DMF/ethanol for all three biomarkers after explicit solubility testing; PBS specifically avoided because its phosphate peaks overlap the 800–1100 cm⁻¹ fingerprint region used for biomarker diagnostic bands |
| Acetone / IPA | Cleaning (alternative dry-processing route) | 5 min each | Used in a considered-but-not-executed alternative (Ag sputtering + anneal) route |

Post-deposition: DI rinse → 110 °C bake, 10 min → vacuum-desiccated storage (current mitigation against silver oxidation; no formal shelf-life study exists yet).

# 4. Biomarker Reagents — Currently In Hand

| Biomarker | Identity / Supplier | Purity | Status / notes |
| --- | --- | --- | --- |
| NAM (N-acetylmuramic acid) | Sigma-Aldrich A3007, CAS 10597-89-4, MW 293.27 g/mol | ≥98% (TLC) | In hand. Melting point 125 °C; storage 2–8 °C; ≈50 mg/mL solubility in water |
| LPS (lipopolysaccharide) | Sigma-Aldrich L2630, E. coli O111:B4, phenol-extracted, CAS 93572-42-0 | Protein impurity ≤3% | In hand (10 mg pack). Procured Dec 2025–Jan 2026; ≈100 experiments at 0.1 mg/experiment; ≈5 mg/mL solubility (hazy), up to ~20 mg/mL warmed |
| LTA (lipoteichoic acid) | Sigma-Aldrich L2515, Staphylococcus aureus | >95% | In hand (10 mg pack). Procured Dec 2025–Jan 2026; ≈100 experiments at 0.05 mg/experiment; highest-cost reagent of the three (₹30,000–40,000); ≈5 mg/mL solubility (hazy) |

This is the full active biomarker panel — all three procured reagents needed for the project's stated Gram-class discrimination goal (NAM = pan-bacterial, LPS = Gram-negative, LTA = Gram-positive) are already in hand. No further biomarker procurement is needed to proceed with detection-chemistry work on any of the three.

## Reference / calibration compounds

| Compound | Role | Status |
| --- | --- | --- |
| Rhodamine 6G (R6G) | Standard SERS benchmark dye — used to validate the substrate is genuinely SERS-active before testing real biomarkers | In use (97 spectra collected across 8 fabrication conditions); supplier/purity/lot not documented — identified gap |
| Methylene blue | Dye reference / contamination check | Used at 1 µg/mL in several NAM and LTA datasets |
| Glucose | Non-biomarker training reference | Used for KnowItAll® software peak-library validation |

# 5. Capture-Chemistry Reagents — NOT Yet Procured

This is the single biggest open decision point for the project right now: none of the surface-functionalisation chemistry needed to move past bare-metal detection is currently in the lab. Every strategy document reviewed treats this as an external dependency, not a data-driven blocker.

| Reagent | Intended target | Status |
| --- | --- | --- |
| 4-Mercaptophenylboronic acid (4-MPBA) | NAM (cis-diol capture); also relevant to LPS core sugars (KDO/heptose) | Not available. Explicitly deferred — "Exp. 4 (deferred until MPBA available)" in the current SOP. This is the literature-prioritised #1 capture chemistry for NAM |
| Cysteamine | Generic cationic/electrostatic capture (NAM carboxylate; LPS phosphates) | Not available |
| Chitosan | Generic cationic capture, LPS-precedented | Not available |
| Polymyxin B | Lipid A-specific capture for LPS | Not available — named in the drafted blood-sample capture workflow, not yet ordered |
| Anti-Lipid A / anti-LPS antibody | Selective LPS capture / reporter (Rusciano-style) | Not available |
| Anti-LTA antibody | Selective LTA capture | Not available — named in the drafted blood-sample capture workflow, not yet ordered |
| Vancomycin | Universal peptidoglycan (NAM-proximal) capture probe | Not available — literature reports ~1000-fold capture enhancement |
| WGA (wheat germ agglutinin)-functionalised gold nanostars | GlcNAc/peptidoglycan capture | Not available |

Decision relevance: the current bare-Ag/Si substrate can be tested against all three biomarkers today with zero further procurement (Section 4), and for LPS specifically this is expected to work reasonably well without any capture layer (see the LPS capture-strategy note). NAM is the one biomarker where bare-substrate detection is not expected to succeed on its own — MPBA is the literature-prioritised next reagent to order if NAM-specific work is the priority.

# 6. Computational Infrastructure

| Component | Detail |
| --- | --- |
| DFT software | Gaussian 09 and Gaussian 16 (ORCA attempted early, abandoned) |
| Compute resource | IIT Jodhpur institute HPC cluster (approved under Prof. Ajay Agarwal, project "Sepsis Biomarkers Analysis using SERS Sensor") |
| Structure-building tools | Avogadro, GaussView; RDKit-based Python scripts for programmatic Gaussian/ORCA input generation |
| Typical job resources | 16 CPU cores, 48 GB memory (NAM final calculation) |
| Spectral ID software | KnowItAll® Raman Identification Pro (adopted Jan 2026, validated on reference glucose spectrum) |

# 7. Substrate Characterisation Instruments — Listed vs. Actually Used

This distinction matters directly for decision-making: several instruments are named as intended methodology but have no completed output on file for this project's own substrate.

| Instrument | Status |
| --- | --- |
| Optical microscope | Listed as intended (Nov 2025 report); no completed characterisation on file for this project's substrate |
| SEM / FESEM | Listed as intended; the only FESEM images on file belong to the host lab's earlier, separate sweat-metabolite sensor publication — not this project's own substrate |
| AFM | Listed as intended; no data on file |
| UV-Vis (plasmon resonance) | Listed as intended, specifically flagged as needed to confirm the fabricated silver's plasmon resonance is matched to the 785 nm excitation line; no data on file |
| Profilometer / contact angle | Listed as intended; no data on file |
| XRD | Not evidenced as used; would confirm metallic (not oxidised) Ag phase |

Fabrication equipment actually evidenced as used: standard wet-chemistry glassware/hot plate, HF-handling equipment, and an oven/hot plate for the 110 °C bake. A sputtering-and-annealing alternative route was drafted in planning notes but no sputterer, evaporator, PECVD, RIE, or spin-coater output was found — this equipment does not appear to have been used.

# 8. Summary — What This Means for the Next Decision

| Question | Answer from current inventory |
| --- | --- |
| Can any biomarker be tested today with zero new procurement? | Yes — NAM, LPS, and LTA are all in hand, and the Ag/Si substrate + 785 nm Raman system is fabricated and R6G-validated |
| Which biomarker is best matched to the current bare substrate? | LPS — amphiphilic and multiply anionic, with direct literature precedent for bare-metal detection using the same E. coli O111:B4 serotype (see the LPS capture-strategy note) |
| Which biomarker needs new chemistry before it can realistically work? | NAM — bare-Ag detection is not expected to succeed; MPBA is the literature-prioritised next reagent, not yet ordered |
| What single measurement would most improve confidence across all three? | True laser power in mW (power meter) and SEM/AFM of the actual fabricated substrate — both cheap, both currently missing, both needed for any quantitative enhancement-factor or cross-paper comparison |
| What is the fastest, lowest-cost next experiment overall? | Bare-substrate LPS drop-cast, since this reagent is a near-exact match to a published benchmark (Rusciano et al. 2023) and needs no new chemistry |