Measurement Worksheet

R6G 3-Point Calibration + NAM Detection Test — Etched Black-Si/Ag Substrate

____________________________________________________________

Sukesh P (M24IM1007)  |  12 August 2026

# 1. Order of Measurements

Piece A, Spot 1 — BLANK: bare etched substrate, no drop. Measure first, before anything touches the surface.

Piece B (2nd sensor), Spot 1 — BLANK repeat: same settings, bare surface. Confirms the blank is reproducible / substrate-to-substrate consistent, not a one-off.

Piece A, Spot 2 — R6G 0.001 mM (1 µM): drop 5-10 µL, air-dry fully, then measure. Lowest concentration first.

Piece A, Spot 3 — R6G 0.01 mM: fresh spot, same drop volume, dry, measure.

Piece A, Spot 4 — R6G 0.1 mM: fresh spot, same drop volume, dry, measure. Highest concentration last.

Piece B, Spot 2 — NAM (top concentration): use the second piece, not Piece A. Keeps NAM spot away from any R6G residue, so a positive result can't be a false one from R6G carryover.

Fresh spot = a location at least a few mm from any earlier drop. Reuse a clean pipette tip per concentration, low to high, to avoid carrying trace R6G into the next drop.

# 2. Raman Settings

## Keep IDENTICAL for every measurement in Step 1 (blank + all 3 R6G points)

Laser wavelength / line — whatever you select first, do not switch mid-series.

Objective / spot size / focus procedure — refocus fresh each time, but same objective.

Spectral range — 400-1800 cm⁻¹ covers every peak you need below.

Laser power and integration time — pick one setting and hold it fixed across blank + all 3 R6G concentrations. This is what makes the concentration series meaningful: rising or falling signal must come from concentration, not from a settings change.

Accumulations — 3 scans per spot, averaged, at every point.

## Starting point

Start at low power (lowest step available, e.g. ~1-10%) with 10 s integration, 3 accumulations. If the blank is clean (flat, no burn mark, no new peaks) and the lowest R6G spot (0.001 mM) still shows no peaks at this setting, step power up in one increment and re-measure — but if you change the setting, re-measure the blank at the new setting too before trusting an R6G result at that setting.

## For NAM (Piece B)

NAM is a much weaker scatterer than R6G, so you are allowed to raise power and/or integration time above whatever you used for R6G. But whatever setting you land on, measure the Piece B blank (Step 2) at that exact same setting — a blank taken at different settings than the sample is not a valid comparison.

⚠ If any spot looks burnt/discolored after a scan, that spectrum is compromised — move to a fresh spot and drop power before continuing.

# 3. R6G — Peaks to Check

These are the standard R6G SERS bands. If your spectrum is real R6G signal (not noise/fluorescence), you should see most of these, strongest ones first:

| cm⁻¹ | Assignment | Priority |
| --- | --- | --- |
| 612 | C-C-C ring in-plane bend | check |
| 774 | C-H out-of-plane bend | PRIMARY — should be clearly visible if any R6G signal is present |
| 1183 | C-H in-plane bend | check |
| 1310 | aromatic C-C stretch | check |
| 1362 | aromatic C-C stretch | PRIMARY |
| 1509 | aromatic C-C stretch | PRIMARY — strongest R6G band in most SERS spectra |
| 1575 | aromatic C-C stretch | check |
| 1650 | aromatic C-C stretch | check |

Expect: peak count/sharpness should hold up (or degrade gracefully) as concentration drops from 0.1 → 0.001 mM. A rough intensity trend across the 3 points is your calibration check that the substrate is SERS-active before you trust the NAM result.

# 4. NAM — Peaks to Check

Pulled from your own reference NAM Raman spectrum (manuscript Table, 22 independent measurements, 100% reproducibility unless noted) — not generic literature values:

| cm⁻¹ | Assignment | Note |
| --- | --- | --- |
| 930 | Ring breathing, ν(C-C-O) | PRIMARY — single strongest band in reference spectrum |
| 871 | γ(CH₂) hydroxymethyl wagging | PRIMARY — 2nd strongest |
| 541 | Skeletal bending, ring torsion | strong |
| 1023 | ν(C-O) + δ(C-O-H) | strong |
| 1347 | δ(C-H) / δs(CH₃) umbrella | strong |
| 1394 | δ(C-H) / δs(CH₃) umbrella | strong |
| 458 / 488 | Skeletal bending, ring torsion | check if low range is scanned |

Do not use ~730-744 cm⁻¹ as a marker band — your own DFT normal-mode analysis showed this mode is dominated by a single H-atom motion (carboxylic O-H out-of-plane), not the ring-breathing mode some bacterial-SERS papers assume. 930 cm⁻¹ is the correct ring-breathing band.

Detecting any 2-3 of the primary/strong bands above, at the correct relative positions, is a legitimate positive signal. A single unassigned bump is not enough.

# 5. Log While You Measure

| Spot | Sample | Power | Integ. time | Peaks seen (list cm⁻¹) | Notes |
| --- | --- | --- | --- | --- | --- |
| A-1 | Blank |  |  |  |  |
| B-1 | Blank (repeat) |  |  |  |  |
| A-2 | R6G 0.001 mM |  |  |  |  |
| A-3 | R6G 0.01 mM |  |  |  |  |
| A-4 | R6G 0.1 mM |  |  |  |  |
| B-2 | NAM (top conc.) |  |  |  |  |