# NAM Raman Study — Data Archive

N-Acetylmuramic Acid: Experimental Raman Spectroscopy and DFT Vibrational Analysis

Sukesh P · Department of Electrical Engineering, Indian Institute of Technology Jodhpur Archive compiled 5 August 2026

## 1. SAMPLE

| Compound | N-Acetylmuramic acid |
| --- | --- |
| Abbreviations | NAM, MurNAc, NAMA |
| Supplier | Sigma-Aldrich (Merck) |
| Catalogue number | A3007-100MG (100 mg pack) |
| CAS number | 10597-89-4 |
| Lot number | BCCK5768 |
| Formula | C₁₁H₁₉NO₈ |
| Molecular weight | 293.27 g/mol (anhydrous — not a hydrate) |
| Purity | ≥98% (TLC) |
| Source | Synthetic · manufactured in Switzerland |
| Form | White powder |
| Melting point | 125 °C |
| Stated impurity | ≤1 mol/mol methanol |
| Storage | 2–8 °C · amber glass bottle |
| Sample preparation | Powder placed directly on clean glass microscope slide, used as received, no purification or recrystallisation |

### Structure

SMILES

CC(O[C@H]1[C@H](O)[C@@H](CO)OC(O)[C@@H]1NC(C)=O)C(O)=O

InChI

1S/C11H19NO8/c1-5(11(18)19)20-10(9(17)8(16)4-14)7(3-13)12-6(2)15/
h3,5,7-10,14,16-17H,4H2,1-2H3,(H,12,15)(H,18,19)/t5-,7+,8-,9-,10-/m1/s1

IMPORTANT — anomeric configuration is NOT specified. The anomeric carbon in the SMILES carries no stereo descriptor, and the InChI stereo layer defines only five centres. The material is an α/β anomeric mixture, normal for a reducing sugar. The DFT calculation models the β anomer only.

## 2. INSTRUMENT

| System | B&W Tek i-Raman Plus |
| --- | --- |
| Model | BWS465-785H |
| Spectrometer | BTC665N-785H-SYS |
| Microscope | BAC102-785E Raman microscope |
| Objective used | 20× Plan |
| Other objective available | 50× LMPlan (NA 0.50, WD 10.5 mm) — not used |
| Laser wavelength | 785 nm (784.92 nm reported by software) |
| Laser power | Safety label: 495 mW max output · Spec: 455 mW at laser port, 340 mW exiting probe · 0–100% in 1% steps |
| Laser class | Class 3B |
| Specified range | 65–2800 cm⁻¹ |
| Detector | 2048-pixel TE-cooled CCD, sensor at −2 °C |
| Spectral range | −46 to 2842 cm⁻¹ |
| Spectral sampling | 1.76 cm⁻¹/pixel at 400 cm⁻¹ → 1.43 cm⁻¹/pixel at 1800 cm⁻¹ |
| Spectral resolution | < 3.5 cm⁻¹ (at 912 nm, atomic emission lines) |
| Power at sample | _______________ mW at 70% ← measure with power meter (upper bound: 70% × 340 mW = 238 mW at probe, less at sample) |
| Dark correction | Applied by instrument software (column Dark Subtracted #1) |

### Nominal power reference

Three different figures apply — do not conflate them:

| Figure | Value | What it means |
| --- | --- | --- |
| Safety label | 495 mW | Maximum accessible emission, for Class 3B hazard classification |
| Spec, laser port | 455 mW | Manufacturer nominal at the port |
| Spec, exiting probe | 340 mW | What leaves the fibre-optic probe |
| At sample | unmeasured | Lower again — through BAC102 microscope + 20× objective |

Based on 340 mW nominal exiting the probe:

| Setting | Exiting probe |
| --- | --- |
| 5% | 17 mW |
| 40% | 136 mW |
| 70% | 238 mW |
| 90% | 306 mW |

These are at the probe, not at the sample. Further losses occur through the BAC102 microscope and the 20× objective, so incident power is lower and still needs direct measurement.

## 3. FILE NAMING

NAM_P70_T25s_A5_R2_03.csv
│   │   │    │  │  └── sequence number within the set
│   │   │    │  └───── set code (see below)
│   │   │    └──────── accumulations
│   │   └───────────── integration time in seconds
│   └───────────────── laser power, percent of nominal
└───────────────────── sample: NAM powder, or GLASS for blank

### Set codes

| Code | Meaning |
| --- | --- |
| GRID | Power × integration optimisation grid (April session) |
| R1–R9 | Replicate sets — fresh spot for each measurement |
| SS1 | Photostability — same spot, consecutive scans |
| DMG1, DMG2 | Sample destroyed by laser — excluded from analysis |
| BLANK | Empty glass slide |

## 4. FOLDER CONTENTS

NAM_Raman_Archive/
├── 00_READ_ME_FIRST.md          this file
├── FILE_INDEX.csv               every file with its conditions
│
├── 01_Raman_Data/               227 spectra
│   ├── A_Optimisation_Grid_April/     98
│   ├── B_Replicates_May/             104
│   ├── C_Photostability_May/          10
│   ├── D_Glass_Blanks/                 5
│   └── E_EXCLUDED_Laser_Damage/       10
│
├── 02_DFT_Calculation/          Gaussian 16 log, input, geometry, all 111 modes
├── 03_Processed_Results/        reference spectrum, peak list, comparison table
├── 04_Figures/                  publication figures + diagnostics
└── 05_Manuscript/               LaTeX, Word, PDF, supplementary

### 01_Raman_Data — what each folder is for

A_Optimisation_Grid_April (98 spectra, 20–24 April) Systematic grid: 5 integration times (5, 10, 15, 20, 25 s) × 16 power levels (5–80% in 5% steps). Produces the acquisition-optimisation figure. One spectrum per condition — these are different conditions, not replicates.

B_Replicates_May (104 spectra, 29–30 May) Nine replicate sets, each at a single fixed condition with a fresh spot per measurement. This is the dataset the reference spectrum is built from. Set R9 was found misfiled inside a folder named “glass slide empty” — those files are NAM powder, not glass.

C_Photostability_May (10 spectra) Ten consecutive scans at 90% / 30 s without moving the stage. Demonstrates sample stability under prolonged illumination at that condition.

D_Glass_Blanks (5 spectra) Genuine empty-slide blanks. Caution: recorded at 5% / 5 s / 1 accumulation only, which does not match the sample measurement conditions. Blanks at 70% / 25 s / 5 accumulations are still needed for the substrate control.

E_EXCLUDED_Laser_Damage (10 spectra) — DO NOT USE AS NAM SPECTRA 90% power with 60 s integration destroys the sample. Correlation with the reference spectrum falls to r = 0.01–0.40, and the 930 cm⁻¹ band decays 0.44 → 0.23 across five consecutive same-spot scans. Retained deliberately: they establish the damage threshold.

## 5. DAMAGE THRESHOLD

| Condition | Outcome |
| --- | --- |
| ≤80% power, ≤25 s | safe |
| 70% power, 30 s | safe (r = 0.977) |
| 70% power, 60 s | safe (r = 0.970) |
| 90% power, 30 s | safe (r = 0.948) |
| 90% power, 60 s | SAMPLE DESTROYED |

## 6. DFT CALCULATION

| Software | Gaussian 16 |
| --- | --- |
| Method | B3LYP with Grimme D3 dispersion, Becke–Johnson damping (D3BJ) |
| Basis set | 6-311++G(d,p) |
| Route | # opt=(calcfc,tight) freq=raman b3lyp/6-311++g(d,p) empiricaldispersion=gd3bj geom=connectivity int=ultrafine scf=(tight,xqc) |
| Phase | Gas phase, isolated molecule (no solvation model) |
| Starting geometry | β anomer, PubChem CID 5462244 |
| Atoms | 39 (C₁₁H₁₉NO₈) |
| Normal modes | 111 |
| Imaginary frequencies | 0 — true minimum confirmed |
| Termination | Normal |
| Frequency scaling | 0.980 below 1800 cm⁻¹, 0.967 above |

## 7. PROCESSING PIPELINE

Applied identically to every spectrum:

- Extract Raman Shift and Dark Subtracted #1

- Interpolate to a uniform 1 cm⁻¹ grid, 400–1800 cm⁻¹

- ALS baseline correction (λ = 1×10⁵, p = 0.001, 10 iterations)

- Savitzky–Golay smoothing (window 11, polynomial order 3)

- Min–max normalisation

- Build reference spectrum from replicate sets R2 + R4 (n = 22); average, compute standard deviation

- Peak detection, validated against local noise on two independent criteria

- Match to calculated modes within ±18 cm⁻¹, preferring highest Raman activity

## 8. KEY RESULTS

- Reference spectrum from 22 independent replicate measurements (sets R2 + R4)

- 41 reproducible bands, 400–1800 cm⁻¹ (32 high confidence, 9 tentative)

- 30 matched to calculated normal modes

- MAE 9.3 cm⁻¹ · RMSE 10.3 cm⁻¹ · r = 0.9996

- The mode at 731 cm⁻¹ is out-of-plane carboxylic acid O–H deformation (47.9% of displacement), with only 0.5% pyranose ring character — arguing against a muramic acid origin for the bacterial 730 cm⁻¹ SERS marker band

- The 872 cm⁻¹ band is assigned to hydroxymethyl wagging (calculated 873 cm⁻¹)

## 9. STILL OUTSTANDING

- ☒ Lot number from the bottle — BCCK5768

- ☐ Power at sample in mW (power meter at objective focal plane)

- ☒ Spectral resolution — < 3.5 cm⁻¹ (manufacturer spec)

- ☐ Glass blanks at 70% / 25 s / 5 accumulations — 3 spots, 10 minutes

- ☒ Extend range above 2842 cm⁻¹ — not possible; BWS465-785H is specified to 2800 cm⁻¹

Note: an earlier partial copy exists at mtp/NAM_Raman_Study_2026. It has filename collisions and is superseded by this archive — safe to delete.