# Measuring Laser Power at the Sample

Update — you do not need this to submit the paper.

The manuscript has been rewritten to report power settings as percentages of the instrument maximum, together with the manufacturer’s nominal figures and an explicitly-labelled estimated range for the power density. That is honest, publishable, and common practice in the Raman literature.

This protocol is kept for if a meter becomes available later, or if a reviewer asks. Section 3 lists free routes that need no equipment.

Instrument: i-Raman Plus BWS465-785H + BAC102-785E microscope, 20× Plan, 785 nm

Why: the paper reports a laser-damage threshold. Stated as “90% power” it is unreproducible on any other instrument. Stated as a power density in mW µm⁻² it is a physical quantity anyone can check.

## 1. What you need

An optical power meter with a sensor rated for 785 nm and at least ~500 mW.

| Type | Suitable? | Note |
| --- | --- | --- |
| Thermal (thermopile) | Best | Handles high power, wavelength-flat, won’t be damaged |
| Photodiode | Usable | Must set wavelength to 785 nm; can be damaged by a focused beam |

Common models: Thorlabs PM100D/PM16 series, Ophir Nova/Vega, Coherent LabMax, Gentec. Any photonics, laser or optics lab at IIT Jodhpur will have one — this is standard equipment, worth asking around before buying.

## 2. Procedure

- Set the meter to 785 nm. Photodiode sensors are strongly wavelength dependent; a wrong setting gives a wrong answer with no warning.

- Zero the meter with the laser off or shuttered.

- Position the sensor at the sample plane — where the powder normally sits, at the objective’s working distance. This is the measurement that matters. Measuring anywhere else gives the wrong number.

- Defocus slightly if using a photodiode. Move the sensor 1–2 mm below the focal plane so the beam has expanded. All the light is still collected, but the intensity on the detector element is far lower. A focused 785 nm beam at >100 mW can permanently damage a photodiode sensor. Thermal sensors are safe at focus.

- Record at several settings — 5%, 40%, 70%, 90%. Four points confirm the response is linear and give you the whole curve rather than one point.

- Let each reading settle, especially with a thermal sensor, which takes several seconds to equilibrate.

Wear the laser safety goggles supplied with the instrument. This is a Class 3B laser and you will have the beam path open.

### Record this

| Setting | Reading (mW) |
| --- | --- |
| 5% | ______ |
| 40% | ______ |
| 70% | ______ |
| 90% | ______ |

## 3. If no power meter is available — free routes

In order of effort. None requires buying anything.

### a. Ask Metrohm — do this first

One email to your Metrohm/B&W Tek India representative:

“For the BWS465-785H with the BAC102-785E microscope and a 20× Plan objective, what is the approximate optical transmission from probe output to sample, or the typical power at the sample at a given percentage setting?”

They may have the figure already. Costs nothing, and an answer from the manufacturer is citable.

### b. Central Instrumentation Facility

IIT Jodhpur’s CIF, and the physics and materials departments, will almost certainly have an optical power meter. Photonics and laser groups use them routinely. A twenty-minute loan is all that is needed.

### c. Any group running a laser

Ask around. Power meters are common shared equipment and people generally lend them without ceremony.

### d. Estimate from component specifications — what the paper now does

A 20× Plan achromat transmits roughly 70–85% at 785 nm; microscope optics (beamsplitter, filters, mirrors) perhaps 50–80%. Combined, about 40–65%.

This is what the manuscript now reports, clearly labelled as an estimate derived from component specifications rather than a measurement. That distinction is what makes it acceptable.

### e. Low-cost meter, if you ever want one

Basic laser power meters covering 785 nm at a few hundred mW are available for roughly ₹5,000–15,000. Accuracy is modest but adequate for an order-of-magnitude figure. Not necessary for this paper.

## 4. Converting to power density

### Spot size

For a diffraction-limited focus:

d = 1.22 λ / NA

With λ = 0.785 µm and a 20× Plan objective (NA ≈ 0.40):

| NA | Spot diameter | Area |
| --- | --- | --- |
| 0.35 | 2.74 µm | 5.88 µm² |
| 0.40 | 2.39 µm | 4.50 µm² |
| 0.45 | 2.13 µm | 3.56 µm² |

Check the NA marked on your objective barrel — it is printed on the side.

Note the real spot is often larger than the diffraction limit, because the probe delivers an image of the fibre core rather than a perfect point. If your BAC102 has a video microscope, you may be able to see the spot and estimate its diameter against a stage micrometer. That is a better number than the theoretical one.

### The calculation

power density = P_sample / spot area

### Expected range for your system

Taking 238 mW at the probe for the 70% setting:

| Transmission | P at sample | Power density |  |
| --- | --- | --- | --- |
| 30% | 71 mW | 15.9 mW µm⁻² | 1.6 MW cm⁻² |
| 40% | 95 mW | 21.1 mW µm⁻² | 2.1 MW cm⁻² |
| 50% | 119 mW | 26.4 mW µm⁻² | 2.6 MW cm⁻² |
| 60% | 143 mW | 31.7 mW µm⁻² | 3.2 MW cm⁻² |
| 70% | 167 mW | 37.0 mW µm⁻² | 3.7 MW cm⁻² |

(1 mW µm⁻² = 0.1 MW cm⁻²)

These numbers make your damage result physically sensible. Power densities of a few MW cm⁻² will readily char an organic powder over 60 seconds, which is exactly what you observed at 90% / 60 s. Percentages conceal that; power density explains it.

## 5. What to report in the paper

Once measured, the Methods section should state:

“Laser power at the sample was measured with a [meter model] at the objective focal plane and was X mW at the [Y]% setting used for the final measurements. With a 20× objective (NA Z), this corresponds to a spot diameter of approximately D µm and a power density of approximately P mW µm⁻².”

And Section 3.3 should restate the damage threshold in those terms:

“Sample degradation occurred at a power density of approximately P mW µm⁻² with 60 s exposure, while P′ mW µm⁻² for 30 s produced no detectable change over ten consecutive scans.”

That is the version another laboratory can act on.

## 6. Send me

- The four meter readings

- The NA printed on your objective

- Spot diameter, if you can measure it

I will compute the power densities, update the Methods and Section 3.3, and rebuild both output formats.