Sigma-Aldrich MurNAc

https://www.sigmaaldrich.com/IN/en/product/sigma/a3007?srsltid=AfmBOor6RMD2491NjmxDPcxMJuOyywBG89HfqEkaSXxflAKH7u3lgWJO


PubChem MurNAc

https://pubchem.ncbi.nlm.nih.gov/compound/5462244#section=Information-Sources

| Item | PubChem | Sigma-Aldrich |
| --- | --- | --- |
| Name | N-Acetylmuramic acid | N-Acetylmuramic acid |
| Formula | C₁₁H₁₉NO₈ | C₁₁H₁₉NO₈ |
| MW | 293.27 g/mol | 293.27 g/mol |
| CAS | 10597-89-4 | 10597-89-4 |
| Structure | MurNAc | MurNAc |
| Stereochemistry | D-MurNAc | D-MurNAc |

Sigma-Aldrich sells the same molecule that PubChem documents.
PubChem = reference database
Sigma-Aldrich = physical supplier

The molecular structure of N-acetylmuramic acid (MurNAc) was obtained from the PubChem database (CID 5462244). The structure corresponds to the commercially available MurNAc standard (Sigma-Aldrich, ≥98% purity)














1️⃣ MATERIAL & SAMPLE REQUIREMENTS (EXPERIMENTAL)

1.1 Chemical standard (already done ✔)

- Compound: N-Acetylmuramic acid (MurNAc)

- Purity: ≥98% (Sigma-Aldrich A3007)

- Form: Powder

- Solubility: Water (up to ~50 mg/mL)

📌 Mention in report:

“Analytical-grade MurNAc (≥98%, Sigma-Aldrich) was used without further purification.”

1.2 Substrate

- CRS / SERS chip

- Nanostructure: Au/Ag nanoparticles (gold–silver nano-coated)

- Role: Localized surface plasmon enhancement

📌 Critical:
You must record:

- Au only / Ag only / Au–Ag hybrid

- Approximate particle size (if vendor gives it)

- Chip batch ID

1.3 Sample preparation (VERY IMPORTANT)

You must define and keep constant:

| Parameter | Requirement |
| --- | --- |
| Solvent | Milli-Q water (preferred) |
| Concentration | Start: 1 mM → dilute (100 µM, 10 µM) |
| Volume on chip | e.g., 5–10 µL |
| Drying | Air-dry / N₂ dry |
| Replicates | ≥3 spots |

📌 Why this matters:
SERS peak intensity changes, but peak positions should remain stable.

2️⃣ RAMAN / SERS INSTRUMENT SETTINGS (LOCK THESE)

You must use identical settings for comparison.

| Parameter | What to report |
| --- | --- |
| Laser wavelength | 532 / 633 / 785 nm |
| Laser power | (e.g., 0.5–5 mW on sample) |
| Objective | 10× / 50× |
| Integration time | (e.g., 5–10 s) |
| Accumulations | 2–5 |
| Spectral range | 400–1800 cm⁻¹ |
| Spectral resolution | cm⁻¹ value |

📌 Golden rule:
Never change settings between samples.

3️⃣ CONTROL EXPERIMENTS (MANDATORY)

You must record all three:

1️⃣ Bare CRS chip (background)
2️⃣ Solvent on chip (water only)
3️⃣ MurNAc on chip

📌 This proves peaks are from MurNAc, not the substrate.

4️⃣ THEORETICAL REQUIREMENTS (COMPUTATIONAL)

4.1 Structure preparation

- Use SDF from PubChem / Sigma

- Convert → XYZ

- Check:

  - Correct stereochemistry

  - Neutral (COOH) form

  - No missing hydrogens

4.2 DFT Raman calculation (minimum acceptable level)

| Item | Recommendation |
| --- | --- |
| Software | ORCA / Gaussian |
| Method | B3LYP |
| Basis set | 6-31G(d) or def2-SVP |
| Job | Geometry optimization + Frequency |
| Raman | YES (mandatory) |
| Scaling factor | ~0.96–0.98 |

📌 You must confirm:
✔ No imaginary frequencies → stable structure

5️⃣ EXPECTED RAMAN-ACTIVE REGIONS (VERY IMPORTANT)

You should expect and later assign:

| Raman shift (cm⁻¹) | Assignment |
| --- | --- |
| 850–950 | C–O–C, sugar ring |
| 1000–1150 | C–O stretching |
| 1250–1350 | C–N, amide III |
| ~1450 | CH₂ / CH bending |
| ~1550–1650 | Amide I (C=O) |
| ~1730 | COOH (may shift / weaken in SERS) |

📌 Key point:
Peak positions ≈ same
Peak intensities ≠ same (SERS effect)

6️⃣ EXPERIMENT vs THEORY COMPARISON (CORE OF YOUR WORK)

6.1 What MUST match

- Peak positions (±10–20 cm⁻¹)

- Functional group assignments

- Presence of amide + sugar modes

6.2 What will differ (and is OK)

| Aspect | Reason |
| --- | --- |
| Intensity | Orientation + plasmon enhancement |
| Missing peaks | Selection rules + surface binding |
| Shifted peaks | Au/Ag–molecule interaction |

📌 This is NOT an error — this is science.

7️⃣ DATA PRESENTATION (FOR REPORT / PAPER)

You must include:

- Experimental Raman spectrum

- Theoretical Raman spectrum (scaled)

- Overlay plot

- Peak assignment table

- Schematic of MurNAc vibrational modes

📌 Strong sentence to use:

“The agreement between calculated and experimental Raman bands confirms the molecular origin of the observed SERS signatures.”

8️⃣ FINAL SCIENTIFIC CLAIM YOU WILL MAKE

You are NOT just measuring Raman.

You are proving:

✅ Identity of MurNAc
✅ Validity of CRS SERS chip
✅ Reliability of DFT Raman prediction
✅ Practical applicability for bacterial biomarkers

9️⃣ COMMON MISTAKES TO AVOID ❌

- Changing laser power mid-experiment

- Comparing unscaled DFT spectra

- Ignoring background subtraction

- Expecting identical intensities

- Not reporting substrate details