Experiment Plan — LPS

Label-Free SERS Detection of Lipopolysaccharide on Bare Ag/Black-Si

____________________________________________________________

Sukesh P (M24IM1007)  |  IIT Jodhpur & AIIMS Jodhpur Joint IDRP  |  Priority Analyte #1

Bare Ag substrate — no capture chemistry required for this protocol

KEY — Why LPS first: it is the ONLY one of your three target analytes with a directly transferable, published, label-free SERS protocol on a silver substrate — including exact sample volume, drying method, acquisition settings, concentration, and a peak assignment table. Everything in Sections 4-7 below is adapted from Wu, Zhao & Zughaier (2021), Biosensors 11(7):234, DOI 10.3390/bios11070234, who detected intrinsic LPS SERS on silver nanorod arrays with no labels and no capture chemistry.

# 1. Objective and Rationale

## 1.1  Objective

Determine whether your bare 2-min-etched Ag/black-Si substrate produces a detectable, reproducible, label-free SERS spectrum from lipopolysaccharide (LPS), using acquisition conditions adapted from a published silver-substrate LPS protocol.

## 1.2  Why LPS is expected to work where NAM did not

| Property | NAM (failed) | LPS (expected to work) |
| --- | --- | --- |
| Ag surface affinity | No thiol, no strong Ag-binding group; amide-blocked N | Lipid A phosphate groups + long acyl chains give real physisorption to Ag |
| Raman scatterers present | Weak, non-resonant sugar-acid only | Long fatty acyl chains (strong C-C, CH2/CH3 modes) + phosphate + saccharide |
| Published SERS spectrum | None found anywhere in literature | Yes — full peak tables for 9+ LPS types on Ag substrates |
| Published LOD on Ag/Au | N/A | Attograms (AuNP); ~10 fM (AgNP/black-P); 10 nmol/mL routinely used |

The mechanism is structural, not luck: lipid A is an amphiphile with six saturated fatty acyl chains on a phosphate-substituted disaccharide. Those acyl chains are strong Raman scatterers AND drive adsorption onto the metal. This is exactly the combination NAM lacks.

# 2A. CONFIRMED REAGENT — Sigma-Aldrich L2630

KEY — Product identity confirmed from the Sigma-Aldrich product page. This resolves items L1 and L2 below.

| Field | Confirmed value |
| --- | --- |
| Product | Lipopolysaccharides from Escherichia coli O111:B4 |
| Catalogue number | L2630 (10 mg pack: L2630-10MG) |
| CAS number | 93572-42-0 |
| EC number | 297-473-0 ¦ MDL: MFCD00164401 |
| Purification | Phenol extraction |
| Form | Lyophilized powder, off-white to faint yellow |
| Purity / impurities | ≤3.00% protein (Lowry-TCA) |
| Solubility | Water: 5 mg/mL (hazy, faint yellow). Cell culture medium: 1 mg/mL. Up to 20 mg/mL in aqueous saline after vortexing and warming to 70-80 °C |
| Storage | 2-8 °C; shipped ambient |
| Indicative price (Bangalore warehouse) | ₹12,773.50 for 10 mg |

## 2A.1  SAFETY — this reagent is acutely toxic

CAUTION — LPS L2630 carries GHS06 (skull and crossbones), signal word DANGER, hazard statement H300 — FATAL IF SWALLOWED, classified Acute Toxicity Category 2 (Oral). This is a materially more hazardous reagent than NAM or R6G. Storage class 6.1A (combustible, acute toxic Cat. 1/2).

| Requirement | Specification |
| --- | --- |
| Required PPE | Eye shields, gloves, N95-type respirator (per supplier) |
| Weighing | Weigh the lyophilized powder in a fume hood or powder-handling enclosure. Never weigh on an open bench — it is a fine, static-prone, acutely toxic powder |
| Precautionary codes | P264 (wash thoroughly after handling) · P270 (do not eat/drink/smoke when using) · P301+P310+P330 (IF SWALLOWED: immediately call a poison centre/doctor; rinse mouth) · P405 (store locked up) · P501 (dispose of contents/container per regulations) |
| Storage | Locked storage at 2-8 °C, clearly labelled |
| Disposal | Do not pour to drain. Collect all LPS-contaminated tips, tubes and substrates as toxic chemical waste |
| Substrate disposal | Substrates carrying dried LPS remain contaminated — collect them, do not leave them on an open bench after measurement |

## 2A.2  Raman-specific caveat — the 3% protein impurity

CAUTION — L2630 is phenol-extracted and specified as ≤3.00% protein. Protein contributes its own Raman bands — notably amide I near 1650-1660 cm-1 and amide III near 1230-1300 cm-1. The 1330 cm-1 LPS marker band sits close to the amide III region. Do not attribute a band in that region to LPS without also confirming the 989 cm-1 carbohydrate band, which protein does not produce.

If protein interference proves to be a problem, Sigma L3024 (same organism, purified by ion-exchange chromatography, TLR-ligand tested) is a higher-purity alternative worth considering for a follow-up confirmation experiment. L2630 is entirely adequate for establishing first detection.

# 2. VALUES THAT MUST BE CONFIRMED BEFORE STARTING

Nothing below is invented. These are unresolved inputs specific to this LPS experiment.

| # | Value | Status | Where to obtain / why it matters |
| --- | --- | --- | --- |
| L1 | LPS source organism and product number | RESOLVED — Sigma L2630, E. coli O111:B4, phenol extracted, CAS 93572-42-0 | See Section 2A |
| L2 | LPS supplied mass and purity grade | RESOLVED — 10 mg pack, ≤3% protein | See Section 2A; note the protein caveat in 2A.2 |
| L3 | Your actual laser power in mW at the sample | STILL UNRESOLVED from previous sessions | CRITICAL for this experiment — the reference protocol used 100 mW at sample. Your 'power level 5/10' is an arbitrary scale. See Section 5.3. |
| L4 | Pyrogen-free / endotoxin-free water availability | Unknown | LPS work should use pyrogen-free water. Ordinary DI water may carry background endotoxin, contaminating a blank. |
| L5 | Vacuum desiccator availability | Unknown | Reference protocol vacuum-dried the 2 µL drop. If unavailable, air-dry is an acceptable substitute but must be logged as a deviation (Section 5.2). |
| L6 | HF stock concentration (carried forward) | Still open | Affects fabrication bath molarity only, not this LPS protocol. |

CAUTION — Item L3 is the single most important one for this experiment. The published protocol used 100 mW at the sample with 10-30 s integration. If your instrument delivers far less (the BWS465-785H class is specified around 30 mW nominal at the probe), you will need LONGER integration to collect comparable signal — see Section 5.3 for how to compensate.

# 3. Materials, Reagents and Equipment

## 3.1  Reagents

| Reagent | Specification | Amount needed | Purpose |
| --- | --- | --- | --- |
| LPS | Sigma L2630, E. coli O111:B4, phenol extracted, CAS 93572-42-0 (CONFIRMED — see Section 2A) | 1 mg of the 10 mg pack is ample for this whole study | Target analyte |
| Pyrogen-free water | Endotoxin-free / LAL-tested, or freshly-drawn high-purity DI | ~20 mL | LPS solvent and dilutions |
| AgNO₃ | ACS grade solid | 0.170 g per 50 mL bath | Substrate fabrication (unchanged, 20 mM) |
| HF | Concentrated stock (conc. to be confirmed, item L6) | 8.6 mL per bath | Substrate fabrication (unchanged) |
| R6G | 0.01 mM working solution | ~1 mL | Substrate SERS-activity control |
| Acetone, IPA | Standard lab grade | ~50 mL each | Si cleaning |

## 3.2  Equipment

Raman: BWS465-785H / BTC665N-785H-SYS, 785 nm — same instrument as all previous work

Balance: analytical balance readable to 0.1 mg (needed for weighing sub-mg LPS quantities)

Pipettes: P2 or P10 micropipette capable of accurate 2 µL delivery, plus P1000 for stock prep

Vortex mixer and bath sonicator: BOTH are required for LPS — see Section 4.2, this is not optional

Vacuum desiccator: preferred for drying (item L5); air-dry acceptable with logging

PP/PTFE labware: for all HF steps

Hydrophobic barrier pen: wax/PAP pen to confine drops on superhydrophilic black-Si

# 4. LPS Solution Preparation — Exact Quantities

## 4.1  The concentration-unit problem (read before weighing)

CAUTION — LPS is a heterogeneous biological macromolecule — it does NOT have a single molecular weight the way NAM (293.27 g/mol) does. Its mass varies by organism, serotype and preparation, typically ~10-20 kDa for smooth LPS. Therefore you CANNOT convert mg/mL to mM reliably. Report LPS concentration in µg/mL or mg/mL only. Any molar figure would be an invented number.

The reference paper reported 10 nmole/mL, but they quantified their LPS specifically by lipid A content using GC-MS — a step you cannot replicate without that instrumentation. Working in mass concentration is the correct, honest approach for your setup.

## 4.2  Stock solution — 1 mg/mL

| Step | Action | Exact quantity | Notes |
| --- | --- | --- | --- |
| 4.2.1 | Weigh LPS into a clean, endotoxin-free microcentrifuge tube | 1.0 mg (record actual mass to 0.1 mg) | Static-prone powder — tare carefully |
| 4.2.2 | Add pyrogen-free water | 1.00 mL | Gives 1 mg/mL stock |
| 4.2.3 | Vortex | 2 minutes, full speed | LPS forms aggregates/micelles — this is essential |
| 4.2.4 | Sonicate in bath sonicator | 15 minutes at room temperature | Breaks up aggregates; the reference protocol explicitly requires sonication before EVERY dilution |
| 4.2.5 | Vortex again briefly | 30 seconds | Immediately before drawing any aliquot |
| 4.2.6 | Store | 4 °C, labelled with date and concentration | Re-vortex + re-sonicate before each subsequent use |

KEY — The vortex + sonicate step is the single most commonly skipped step in LPS work and the most common cause of irreproducible results. LPS self-assembles into micelles and aggregates in water; without dispersion, the 2 µL you pipette may contain wildly different amounts of LPS from draw to draw. Sonicate before EVERY dilution and before EVERY deposition.

### Expected appearance — haziness is NORMAL, do not filter

Sigma's own preparation note states plainly: LPS forms micelles in EVERY solvent, and hazy solutions are expected in water and in phosphate-buffered saline. Water yields a homogeneously hazy solution. Organic solvents do NOT give clearer solutions — methanol produces a turbid suspension with visible floaters. Therefore:

A hazy, faint-yellow 1 mg/mL solution is a correctly prepared solution. Do not filter it, do not centrifuge it clear, and do not discard it as 'undissolved'.

Do not switch to an organic solvent hoping for clarity — it makes dispersion worse and would change the drying behaviour on the substrate.

Water is the correct solvent for this work, and 1 mg/mL sits well below the 5 mg/mL water solubility ceiling.

Sigma notes that up to 20 mg/mL can be reached in aqueous saline with vortexing and warming to 70-80 °C. That is NOT needed here and is best avoided — heating risks altering the LPS, and saline would leave crystalline salt residue on the substrate after drying, which would add its own spurious Raman features.

## 4.3  Working dilution series

Prepare fresh on the day of measurement from the 1 mg/mL stock. Serial dilution in pyrogen-free water:

| Level | Target conc. | Preparation | Final volume | Purpose |
| --- | --- | --- | --- | --- |
| W1 | 100 µg/mL | 100 µL stock + 900 µL water | 1.00 mL | Primary working concentration — start here |
| W2 | 10 µg/mL | 100 µL of W1 + 900 µL water | 1.00 mL | Sensitivity step 1 |
| W3 | 1 µg/mL | 100 µL of W2 + 900 µL water | 1.00 mL | Sensitivity step 2 |
| W4 | 0.1 µg/mL | 100 µL of W3 + 900 µL water | 1.00 mL | Sensitivity step 3 / approaching literature-relevant trace levels |

Vortex 30 s and sonicate 5 min at each dilution step, not just at the stock stage. Use a fresh pipette tip for every transfer, going low-to-high when depositing.

# 5. Experimental Procedure — Step by Step

## 5.1  Substrate fabrication (Day 1)

Fabricate FRESH substrates — do not use the aged/reused 20 mM bath for this experiment, since LPS is your priority target and deserves a clean, characterized substrate.

| Step | Action | Exact quantity / timing |
| --- | --- | --- |
| 5.1.1 | Solvent clean Si pieces: acetone → IPA → DI water | Sonicate 5 min in each |
| 5.1.2 | Native oxide strip in dilute HF | 4.1 mL 49% HF + 15.9 mL DI = 20 mL (~10% HF); dip 30-60 s |
| 5.1.3 | Rinse in DI, proceed IMMEDIATELY to etch | Within 5 minutes of oxide strip |
| 5.1.4 | Dissolve AgNO₃ in DI | 0.170 g in ~20 mL DI (20 mM final) |
| 5.1.5 | Add concentrated HF to the AgNO₃ solution | 8.6 mL |
| 5.1.6 | Top up with DI to final volume | to 50.0 mL total |
| 5.1.7 | Immerse substrate, start timer | Etch exactly 2 min 00 s |
| 5.1.8 | Remove, rinse thoroughly in DI | ≥30 s running DI rinse |
| 5.1.9 | Dry | N₂ blow-dry, or air-dry in covered container |
| 5.1.10 | Fabricate 2 pieces minimum | One for LPS, one spare/repeat |

CAUTION — Full HF safety protocol: double gloves with HF-rated outer layer, face shield, fume hood, calcium gluconate gel present and in date, never work alone. Waste to the silver/heavy-metal stream, not general acid drain.

## 5.2  LPS deposition (Day 2 — measurement day)

| Step | Action | Exact quantity / timing |
| --- | --- | --- |
| 5.2.1 | Draw hydrophobic barrier rings on the substrate | 3-5 rings, ~3-4 mm apart; let dry fully (~5 min) |
| 5.2.2 | Re-vortex and re-sonicate the LPS working solution | 30 s vortex + 5 min sonication |
| 5.2.3 | Take the SAME-SESSION BLANK spectrum FIRST | Before any liquid touches the substrate — see 5.3 |
| 5.2.4 | Deposit LPS inside a barrier ring | 2.0 µL per spot (matches the published protocol exactly) |
| 5.2.5 | Dry | Vacuum desiccator until visibly dry (preferred), OR air-dry undisturbed |
| 5.2.6 | Record actual dry time | Log in minutes — this is a real variable |
| 5.2.7 | Repeat for each spot | Minimum 9 spots at the primary concentration (see 5.4) |

2.0 µL is taken directly from the reference protocol, not chosen arbitrarily. It is also small enough to limit the spreading/coffee-ring problem you hit on 12 Aug with 3 µL on superhydrophilic black-Si — the barrier ring is the additional safeguard.

## 5.3  Raman acquisition settings

| Parameter | Reference paper | Your setting | Justification |
| --- | --- | --- | --- |
| Wavelength | 785 nm | 785 nm | Identical — no change needed |
| Spectral range | 400-1800 cm⁻¹ | 400-1800 cm⁻¹ | Identical; covers all 7 LPS marker bands |
| Laser power at sample | 100 mW | UNKNOWN (item L3) — use your power level 10 | Your instrument class is ~30 mW nominal; you likely have LESS power than the reference |
| Integration time | 10-30 s | START AT 30 s | Compensates for probable lower laser power. If saturated, step down to 15 s then 10 s |
| Accumulations | not stated | 3 | Consistent with all your prior work |
| Spots per sample | 9 randomly selected | 9 minimum | Matches the published reproducibility protocol exactly |

KEY — Power/time trade-off: your own raw-Si calibration (13 Aug) showed SNR grows FASTER than the sqrt(time) shot-noise law in this instrument — going 5 s → 25 s gave a 5.9x SNR gain, not the 2.2x pure shot noise predicts. This means longer integration is an unusually efficient lever on your specific system, and is the right way to compensate for having less laser power than the reference protocol.

CAUTION — Check for burning at 30 s before committing to the full 9-spot run. Take one test spectrum, then visually inspect the spot under the microscope. LPS is an organic biomolecule and CAN photodegrade. If any discolouration appears, drop to 15 s and re-test.

## 5.4  Full measurement matrix

| Order | Sample | Spots | Settings | Purpose |
| --- | --- | --- | --- | --- |
| 1 | Blank (bare etched substrate) | 3 | Same as below | Same-session background — mandatory |
| 2 | R6G 0.01 mM | 3 | Same as below | SERS-activity gate — confirms substrate is alive BEFORE trusting LPS |
| 3 | LPS 100 µg/mL (W1) | 9 | 785 nm, 30 s, 3 acc., power 10 | Primary detection attempt |
| 4 | LPS 10 µg/mL (W2) | 5 | Identical | Concentration response |
| 5 | LPS 1 µg/mL (W3) | 5 | Identical | Concentration response |
| 6 | LPS 0.1 µg/mL (W4) | 5 | Identical | Approaching LOD |
| 7 | Blank repeat (2nd substrate) | 3 | Identical | Confirms blank reproducibility |

Total: approximately 33 spectra. At 30 s integration × 3 accumulations plus repositioning, budget roughly 2.5-3 hours of instrument time.

# 6. LPS Peak Targets

From Wu, Zhao & Zughaier (2021), Table 1 — SERS peak assignments for Neisseria meningitidis LPS structures measured on a silver nanorod substrate. These are published, measured SERS positions on silver, not calculated or inferred values:

| Δν (cm⁻¹) | Vibrational assignment | Priority |
| --- | --- | --- |
| 550 | β(CH₂) in ring | Check |
| 680 | δ(C-O-C); fatty acid | Check |
| 850 | ν(C-O-C); saccharides (1,4-glycosidic link) | Check |
| 988-989 | β(CH); carbohydrates | PRIMARY — reported as a significant peak for LPS and KDO2-lipid A |
| 1131 | ν(C-C); fatty acid | Check |
| 1330 | δ(CH); phospholipid | PRIMARY — reported as a significant peak, strongest discriminator |
| 1450 | α(CH₃/CH₂), β(CH₃/CH₂) | Check |

KEY — Focus on 989 and 1330 cm⁻1. The paper explicitly identifies these two as the significant peaks in wild-type LPS spectra, with 989 attributed to carbohydrate and 1330 to phospholipid δ(CH). If your substrate produces LPS signal at all, these two should appear first.

Additional positions reported for enteric LPS (E. coli, Salmonella, Vibrio) in the same paper: 543, 552, 981 cm⁻¹ — close variants of the above, differing slightly by organism. Note that different LPS sources give measurably different spectra; this is a genuine feature (it enables bacterial differentiation) but it means you should compare your results against the organism you actually use.

NOT CONFIRMED — Your 12 Aug instrument calibration showed a systematic ~2-3 cm-1 offset (Si TO phonon found at 517.8 rather than the textbook 520-521 cm-1). Apply the same tolerance when matching these LPS positions — accept a peak within roughly ±5 cm-1 of the target.

# 7. Data Analysis and Detection Criteria

## 7.1  Mandatory processing step — despiking

CAUTION — ADD A DESPIKE STEP BEFORE ANY PEAK ANALYSIS. Analysis of your 14 Aug LTA data produced an apparent SNR~7-9 'detection' at 1095 cm-1 that turned out to be a SINGLE-PIXEL detector artifact at pixel 640, recurring at the identical pixel across three consecutive spectra. A real Raman band spans 4-5 pixels on this instrument (as your Si TO phonon does). Any candidate peak occupying only 1-2 pixels must be rejected as an artifact.

Practical rule: reject any feature narrower than 3 pixels (~5 cm⁻¹ on this instrument). Apply a median filter or explicit spike-detection pass to every spectrum before peak fitting.

## 7.2  Processing pipeline

Despike. Median-filter or spike-reject; verify pixel width of every candidate peak.

Baseline correction. Asymmetric least-squares (λ=10⁵, p=0.001) — the reference paper also baseline-corrected, as LPS spectra carry a broad fluorescence background.

Average within a sample. The reference protocol averaged 9 spectra per sample after baseline correction. Do the same, but ALSO evaluate spots individually first (7.3).

Do not normalize across concentrations. Raw intensity differences between W1-W4 are the concentration signal you are looking for.

## 7.3  Detection criteria

| Verdict | Requirement |
| --- | --- |
| DETECTED | Both primary bands (989 and 1330 cm⁻¹) at SNR ≥ 3, each spanning ≥3 pixels, reproducible in ≥5 of 9 spots, at consistent position (±5 cm⁻¹) |
| POSSIBLY DETECTED | One primary band at SNR ≥ 3 across ≥3 of 9 spots, OR both bands at SNR 2-3 — repeat before reporting |
| NOT DETECTED | Neither primary band clears SNR ≥ 3 reproducibly, or all candidates fail the pixel-width test |

# 8. File Naming and Logging

Format: EXP-LPS-{conc}-{TYPE}-S{spot}

| Example | Meaning |
| --- | --- |
| EXP-LPS-AG20-BLANK-S01 | 20 mM AgNO₃ substrate, blank, spot 1 |
| EXP-LPS-AG20-R6G-S01 | R6G activity control, spot 1 |
| EXP-LPS-100UGML-S01 | LPS 100 µg/mL, spot 1 |
| EXP-LPS-10UGML-S03 | LPS 10 µg/mL, spot 3 |
| EXP-LPS-0P1UGML-S05 | LPS 0.1 µg/mL, spot 5 |

Log columns (one row per spectrum): Date | Experiment ID | Substrate ID | AgNO₃ conc | AgNO₃ mass | HF vol | Bath vol | Etch time | Si batch | LPS source/lot | LPS stock conc | LPS working conc | Drop volume | Vortex/sonication done (Y/N) | Dry method | Dry time | Laser power level | Integration | Accumulations | Spot # | Sample type | File name | 989 peak (pos/height/SNR/px width) | 1330 peak (pos/height/SNR/px width) | Notes

# 9. Expected Outcomes and Decision Tree

| Outcome | Interpretation | Next step |
| --- | --- | --- |
| Both 989 and 1330 detected at 100 µg/mL, reproducible | LPS detection works on your bare Ag/black-Si — the project has its first real analyte result | Proceed down the dilution series to establish LOD; then repeat on independent substrates for reproducibility |
| Detected only at the highest concentration | Works but sensitivity-limited | Increase integration time further; consider whether the AgNO₃ optimization series would raise enhancement |
| Peaks present but only in some spots | Deposit inhomogeneity (coffee-ring), not necessarily a substrate failure | Improve drop confinement; map across the ring rather than centre only |
| R6G strong, LPS absent at all concentrations | Substrate is SERS-active but LPS is not adsorbing/enhancing on YOUR substrate morphology | This is the informative negative — proceed to the AgNO₃ optimization series with LPS (not NAM) as the test analyte |
| R6G also weak | Substrate fabrication problem, not an analyte problem | Re-fabricate; do not interpret the LPS result at all |

# 10. Position in the Overall Roadmap

This experiment replaces NAM as the immediate next step, on the reasoning that a working detection on ANY of the three target biomarkers establishes the platform and de-risks everything downstream. Once LPS detection is established:

Use LPS as the test analyte for the AgNO₃ concentration optimization series (5/10/20/40/80 mM) — a working analyte makes substrate optimization interpretable in a way an undetectable analyte never can.

Apply the optimized substrate to LTA (Priority #2).

Return to NAM with the optimized substrate plus MPBA capture chemistry once available (Priority #3).

NAM remains the most novel scientific contribution — no published NAM SERS spectrum exists — but it is the hardest target and should not gate progress on the platform.