Experiment Plan — LTA

Label-Free SERS Detection of Lipoteichoic Acid on Ag/Black-Si

____________________________________________________________

Sukesh P (M24IM1007)  |  IIT Jodhpur & AIIMS Jodhpur Joint IDRP  |  Priority Analyte #2

Gram-positive sepsis marker — companion to Experiment_Plan_LPS.docx

KEY — Status of LTA vs LPS: LTA has WEAKER published SERS precedent than LPS. It appears in the literature mainly as a control/reference material in bacterial endotoxin SERS studies rather than as a purpose-built detection target. There is no published, complete, label-free LTA SERS protocol with a full peak assignment table equivalent to the LPS one. This document therefore combines (a) confirmed reagent data from the Sigma product sheet, (b) confirmed structural chemistry, (c) literature-confirmed vibrational assignments for the functional groups LTA contains, and (d) clearly-labelled predictions derived from that structure.

CAUTION — CORRECTION TO EARLIER WORK — your 14 Aug LTA run was initially analysed as showing a signal at 1095 cm-1 (SNR 7-9 in three spots). On pixel-level inspection this proved to be a SINGLE-PIXEL detector artifact at pixel 640, recurring at the identical pixel across three consecutive spectra. A genuine Raman band on this instrument spans 4-5 pixels (as your Si TO phonon does). The correct verdict for 14 Aug is NOT DETECTED. Section 8.1 makes despiking mandatory before any future peak call.

# 1. Analyte Chemistry — What We Are Trying to Detect

## 1.1  Confirmed structure (Sigma product sheet, L2515)

LTA is a linear polymer of phosphodiester-linked glycerol phosphate, covalently bound to a lipid. The C2 position on the glycerol-phosphate is usually glycosylated and/or D-alanylated. Because of the negatively charged glycerol-phosphate backbone and the hydrophobic lipid tail, the molecule is amphipathic.

## 1.2  Why this structure matters for SERS

| Structural feature | Raman consequence | Ag-surface consequence |
| --- | --- | --- |
| Phosphodiester backbone (PO₂⁻) | Strong, characteristic symmetric stretch near 1080 cm⁻¹ — the single best candidate marker | Negatively charged — can interact electrostatically with the Ag surface |
| Lipid tail (hydrophobic) | CH₂/CH₃ deformation modes near 1450 cm⁻¹; C-C stretches near 1130 cm⁻¹ | Drives physisorption onto the metal, similar to lipid A in LPS |
| Glycerol / glycosyl groups | C-O and C-O-C stretches in the 800-1000 cm⁻¹ region | Polar, hydrophilic |
| D-alanyl ester substituents | Ester C=O near 1730 cm⁻¹; possible weak amide-like features | Adds positive charge, partially offsetting backbone charge |

KEY — LTA sits BETWEEN NAM and LPS in expected difficulty. Like LPS it is amphipathic with a lipid anchor (good for adsorption and for CH2/CH3 signal). Unlike LPS it has a simpler, more repetitive backbone with fewer distinct strong scatterers. Expect a weaker, less feature-rich spectrum than LPS, but a genuinely better prospect than NAM.

# 2. Expected Raman / SERS Peaks

## 2.1  Evidence tiers

Every band below is labelled by how well it is supported. Nothing here is presented as an established LTA reference spectrum, because no such published spectrum was located.

| Δν (cm⁻¹) | Assignment | Evidence tier | Basis |
| --- | --- | --- | --- |
| 1080-1095 | ν_s(PO₂⁻) symmetric phosphate stretch | LITERATURE-CONFIRMED for the functional group | Confirmed assignment: the ~1080 cm⁻¹ band is the symmetric stretch of PO₂⁻ present in nucleic acids, LPS AND teichoic acids. A 1095 cm⁻¹ fit to S. aureus LTA is reported in SERS endotoxin work. |
| 1225-1250 | ν_as(PO₂⁻) antisymmetric phosphate stretch | LITERATURE-CONFIRMED for the functional group | Reported as antisymmetric PO₂⁻ stretch in phosphate/phosphodiesters |
| 1450 | α/β(CH₃/CH₂) lipid deformation | LITERATURE-CONFIRMED (LPS Table 1) | Same lipid deformation band assigned in the published LPS SERS table — LTA also has a lipid tail |
| 1131 | ν(C-C) fatty acid | LITERATURE-CONFIRMED (LPS Table 1) | Same basis as above |
| 850 | ν(C-O-C) saccharide | LITERATURE-CONFIRMED (LPS Table 1) | Applies to the glycosyl substituents on LTA's C2 position |
| ~1730 | ν(C=O) ester | PREDICTION from structure | D-alanyl ester groups; not confirmed in any LTA spectrum |
| 1160, 1120 | phosphate region features | LITERATURE-MENTIONED | Listed among major phosphate absorptions |

## 2.2  CRITICAL specificity limitation

CAUTION — The 1080-1095 cm-1 PO2- band is NOT specific to LTA. The same band appears in nucleic acids, in LPS, and in wall teichoic acid. In whole-bacteria Raman studies, 1095 cm-1 in S. aureus is routinely assigned to NUCLEIC ACID, not to LTA. Therefore: detecting a band at 1080-1095 cm-1 from a purified LTA sample on a clean substrate is meaningful, but that same band could never be used on its own to identify LTA in a real biological sample. Any future clinical claim needs a multi-band pattern, not this one peak.

Practical consequence for this experiment: because your samples are purified LTA on a clean, blank-referenced substrate, a 1080-1095 cm⁻¹ band that passes the artifact tests IS interpretable as LTA-derived. But it must be accompanied by at least one lipid-region band (1450 and/or 1131 cm⁻¹) before being called a detection — a lone phosphate band is too ambiguous.

# 2A. CONFIRMED REAGENT — Sigma-Aldrich L2515

KEY — Product identity confirmed from the Sigma-Aldrich product page. This resolves items T1 and T2 below.

| Field | Confirmed value |
| --- | --- |
| Product | Lipoteichoic acid from Staphylococcus aureus (bacterial cell wall polymer) |
| Catalogue number | L2515 (10 mg pack: L2515-10MG; also 5 mg and 25 mg) |
| CAS number | 56411-57-5 ¦ MDL: MFCD00131526 |
| Form | Lyophilized powder |
| Functional group / lipid type | Phospholipid / polymerizable lipids |
| Solubility | Water, 5 mg/mL — hazy, colourless to faint yellow |
| Storage (solid) | 2-8 °C; shipped ambient |
| Storage (aqueous) | 2 days at 4 °C, OR 2 months as single-use frozen aliquots at -20 °C |
| Biological role | TLR2 (and TLR2/6) agonist; PAMP recognised by Toll-like receptor 2, activating NF-κB |
| Indicative price (Bangalore warehouse) | ₹20,664.93 for 5 mg; ₹36,620.98 for 10 mg |

## 2A.1  Safety and disposal

| Requirement | Specification |
| --- | --- |
| Hazard classification | No GHS pictogram assigned (unlike LPS, which is GHS06 acutely toxic). Storage class 11 — combustible solids |
| Water hazard class | WGK 3 — SEVERELY hazardous to water. This governs disposal, not handling |
| Required PPE | Eye shields, gloves, N95-type respirator (per supplier) |
| Disposal | Do NOT pour LTA solutions to drain — WGK 3 rating requires collection as chemical waste. Includes rinse water from LTA-contaminated glassware |
| Handling | Weigh in a fume hood or enclosure to control the fine lyophilized powder |

LTA is markedly less acutely toxic than LPS (which is H300, fatal if swallowed) — but its WGK 3 water-hazard rating means disposal discipline matters more. Handle LTA waste as carefully as you would the silver/HF streams.

## 2A.2  A note on source organism — relevant to literature comparison

Sigma also supplies LTA from Bacillus subtilis (L3265). This matters because the SERS endotoxin literature that mentions LTA used B. subtilis LTA as its control material, whereas the 1095 cm⁻¹ fit was reported for S. aureus LTA. Sigma explicitly notes that LTA differs between species of Gram-positive bacteria.

Recommendation: proceed with S. aureus L2515. It is the clinically relevant Gram-positive sepsis organism for your project, and it is the source for which a 1095 cm⁻¹ SERS fit has been reported.

Note for later: if you ever need to compare directly against the published endotoxin-SERS control spectra, B. subtilis LTA (L3265) would be the matched material. Not needed for first detection.

# 3. VALUES THAT MUST BE CONFIRMED BEFORE STARTING

| # | Value | Status | Where to obtain / why it matters |
| --- | --- | --- | --- |
| T1 | Which LTA you have / will order | RESOLVED — Sigma L2515, S. aureus, CAS 56411-57-5 | See Section 2A. Still log the LOT number of your actual bottle. |
| T2 | LTA mass available | RESOLVED — 5, 10 or 25 mg pack sizes; 5 mg is sufficient for this protocol | See Section 2A |
| T3 | Laser power in mW at sample | STILL UNRESOLVED | Same open item as every previous experiment. Needed to report absolute conditions and to compare with any literature. |
| T4 | Whether your 14 Aug LTA stock is still usable | Unknown — depends on preparation date and storage | Sigma states aqueous LTA is stable only 2 DAYS at 4 °C, or 2 months frozen as single-use aliquots. If your 14 Aug solution was refrigerated and is now days old, DISCARD IT and prepare fresh. |
| T5 | HF stock concentration | Still open (carried forward) | Fabrication bath molarity only |

KEY — Item T4 may partly explain the 14 Aug null result. Sigma's stability data is strict: 2 days at 4 °C. If your LTA solution had been sitting longer than that, degradation/aggregation is a plausible contributing factor alongside the artifact issue.

# 4. Chemicals, Reagents and Sourcing

## 4.1  Full chemical list for this experiment

| Chemical | Grade / spec | Supplier & product | Qty needed | Purpose |
| --- | --- | --- | --- | --- |
| Lipoteichoic acid | from S. aureus, CAS 56411-57-5, lyophilized powder (CONFIRMED — Section 2A) | Sigma L2515-5MG or -10MG | 5 mg sufficient | Target analyte |
| Silver nitrate (AgNO₃) | ACS reagent ≥99% | Sigma 209139 or equivalent | ~1 g total | MACE catalyst / Ag source |
| Hydrofluoric acid (HF) | 48-51% ACS | Any licensed supplier | ~50 mL | Si etchant + oxide strip |
| Rhodamine 6G | Dye content ≥95% | Sigma 252433 | <10 mg | SERS activity control |
| Deionized water | Type I, 18.2 MΩ·cm | In-house | ~500 mL | Solvent, rinsing |
| Acetone | ACS reagent | Standard | ~100 mL | Si cleaning |
| 2-Propanol (IPA) | ACS reagent | Standard | ~100 mL | Si cleaning, storage |
| Nitrogen gas | Dry, filtered | In-house | — | Drying |
| Calcium gluconate gel 2.5% | HF first-aid | Pharmacy/safety supplier | 1 tube, in date | MANDATORY HF safety |

## 4.2  OPTIONAL substrate-chemistry additions (evaluate, do not assume)

You asked whether the substrate should be made with different chemicals. Here is an evaluation of the realistic options, with a recommendation:

| Option | What it would do | Verdict for THIS experiment |
| --- | --- | --- |
| Keep bare Ag (current) | No added chemistry; relies on LTA's lipid tail physisorbing to Ag | RECOMMENDED — LTA is amphipathic and has a genuine adsorption mechanism, unlike NAM. Test this before adding complexity. |
| Cysteamine SAM (2-aminoethanethiol) | Thiol binds Ag, leaves a positively charged -NH₃⁺ surface | WORTH CONSIDERING LATER — LTA's backbone is negatively charged (phosphodiester), so a positively charged surface should electrostatically attract it. This is a more rational pairing for LTA than for NAM. Sigma M6500. Do NOT combine with the bare-Ag run; make it a separate arm. |
| Poly-L-lysine coating | Positively charged polymer layer | Similar rationale to cysteamine but a thicker layer, which pushes the analyte further from the hotspot — SERS enhancement falls off very steeply with distance. Lower priority. |
| MPBA / boronic acid | Binds cis-diols | NOT APPLICABLE NOW (unavailable), and LTA's glycerol backbone is a less ideal boronate target than a true cis-diol sugar. Lower priority for LTA than for NAM. |
| Persulfate (PDS) in the bath | Strong oxidant | NOT RECOMMENDED — unchanged from the earlier evaluation: no literature precedent in one-step Ag/HF MACE, and it could re-oxidise deposited Ag⁰. |

KEY — Recommendation: run Experiment 1 on BARE Ag first. Only if that fails should you order cysteamine and run the charged-surface arm. Adding surface chemistry before establishing the bare-Ag baseline would repeat the mistake of changing two things at once.

# 5. LTA Solution Preparation — Exact Quantities

## 5.1  Solubility and stability (confirmed, Sigma L2515 product sheet)

Solubility: water, 5 mg/mL, yielding a HAZY colourless to faint-yellow solution. Haziness is normal and expected — it is not undissolved material to be filtered out.

Storage of solid: 2-8 °C.

Storage of aqueous solution: stable 2 DAYS at 4 °C, or 2 MONTHS as single-use frozen aliquots at -20 °C.

CAUTION — The 2-day aqueous stability limit is the tightest handling constraint in this entire protocol. Plan to prepare LTA solution on the day of measurement, or to freeze single-use aliquots immediately after preparation. Never re-freeze a thawed aliquot.

## 5.2  Stock solution — 1 mg/mL

| Step | Action | Exact quantity | Timing / notes |
| --- | --- | --- | --- |
| 5.2.1 | Equilibrate the vial to room temperature before opening | — | ~15 min; prevents condensation into hygroscopic solid |
| 5.2.2 | Weigh LTA into a clean microcentrifuge tube | 1.0 mg (record actual to 0.1 mg) | Analytical balance |
| 5.2.3 | Add DI water | 1.00 mL | Gives 1 mg/mL — well below the 5 mg/mL solubility ceiling |
| 5.2.4 | Vortex | 2 min, full speed | LTA is amphipathic and forms micelles — dispersion is essential |
| 5.2.5 | Sonicate (bath) | 15 min, room temperature | Same rationale as LPS; breaks aggregates |
| 5.2.6 | Inspect | Should be HAZY, colourless to faint yellow | Per the product sheet — do not filter the haze out |
| 5.2.7 | Aliquot | 10 × 100 µL single-use aliquots | Freeze at -20 °C if not used same day |
| 5.2.8 | Vortex 30 s immediately before every use | — | Non-negotiable |

## 5.3  Working dilution series

| Level | Concentration | Preparation | Purpose |
| --- | --- | --- | --- |
| T-W1 | 500 µg/mL | 500 µL stock + 500 µL DI | HIGH — maximises chance of first detection |
| T-W2 | 100 µg/mL | 100 µL stock + 900 µL DI | Primary working concentration |
| T-W3 | 10 µg/mL | 100 µL of T-W2 + 900 µL DI | Sensitivity step |
| T-W4 | 1 µg/mL | 100 µL of T-W3 + 900 µL DI | Sensitivity step / approaching LOD |

Note the higher starting concentration (500 µg/mL) compared to the LPS plan. This is deliberate: LTA has weaker published precedent, so the first objective is simply to establish whether ANY signal is obtainable. Sensitivity refinement comes second.

NOT CONFIRMED — LTA, like LPS, is a heterogeneous polymer with no single reliable molecular weight (chain length varies). Report concentrations in µg/mL only. Do not compute molarity — it would be an invented number.

# 6. Substrate Fabrication and Cleaning

## 6.1  Silicon cleaning (identical to all previous runs)

| Step | Action | Exact quantity / timing |
| --- | --- | --- |
| 6.1.1 | Cut Si pieces from the same wafer batch | ~1 × 1 cm; record wafer ID |
| 6.1.2 | Acetone sonication | 5 min |
| 6.1.3 | IPA sonication | 5 min |
| 6.1.4 | DI water rinse + sonication | 5 min |
| 6.1.5 | N₂ dry | until visibly dry |
| 6.1.6 | Native oxide strip: dilute HF | 4.1 mL 49% HF + 15.9 mL DI = 20 mL (~10% HF); dip 30-60 s |
| 6.1.7 | Quick DI rinse, proceed IMMEDIATELY to etch | within 5 min — oxide regrows |

Piranha (Stage A2) is NOT used, consistent with your recent runs. Log this explicitly so it is not later assumed to have been done.

## 6.2  MACE etch — 20 mM AgNO₃, 2 min (established best condition)

| Step | Action | Exact quantity |
| --- | --- | --- |
| 6.2.1 | Dissolve AgNO₃ in DI | 0.170 g in ~20 mL DI |
| 6.2.2 | Add concentrated HF to the AgNO₃ solution | 8.6 mL |
| 6.2.3 | Top to final volume with DI | to 50.0 mL |
| 6.2.4 | Verify final composition | 20 mM AgNO₃; HF ≈4.8 M if stock is 49% (unconfirmed) |
| 6.2.5 | Immerse substrate; start timer | exactly 2 min 00 s |
| 6.2.6 | Remove and rinse | ≥30 s running DI |
| 6.2.7 | Dry | N₂ blow-dry |
| 6.2.8 | Number of substrates | 3 (LTA main, LTA repeat, spare) |

CAUTION — HF SAFETY — applies to steps 6.1.6 and all of 6.2: double gloves with HF-rated outer layer, face shield, fume hood, calcium gluconate gel present and in date, never work alone. PP/PTFE labware only. Waste to the silver/heavy-metal stream. Use a FRESH bath — do not reuse an aged, already-reacted bath for a priority measurement.

## 6.3  Storage between fabrication and measurement

Measure the same day if at all possible — freshly etched Ag surfaces tarnish/oxidise over time, reducing SERS activity.

If storage is unavoidable: IPA immersion in a labelled, sealed vial, dark, and measure within 24 h.

Log the fabrication-to-measurement interval for every substrate. This is a real variable that has not been tracked so far.

# 7. Experiment Strategies

## 7.1  EXPERIMENT 1 — Bare-Ag LTA detection (do this first)

| Field | Specification |
| --- | --- |
| Question | Does purified LTA give any detectable SERS signal on bare 2-min-etched Ag/black-Si? |
| Variable | LTA concentration (500 / 100 / 10 / 1 µg/mL) |
| Constants | 20 mM AgNO₃, 2-min etch, bare Ag, 785 nm, 30 s, 3 acc., fixed power |
| Substrates | 1 primary + 1 repeat |
| Success criterion | ≥1 phosphate band (1080-1095 or 1225-1250) PLUS ≥1 lipid band (1450 or 1131) at SNR ≥3, ≥3 pixels wide, in ≥3 of 9 spots |
| Duration | 1 day fabrication + 1 day measurement |

## 7.2  EXPERIMENT 2 — Conditional branches

| If Experiment 1 gives… | Then Experiment 2 is… | Rationale |
| --- | --- | --- |
| Detection at 500 and 100 µg/mL | Extend the dilution series downward (0.1, 0.01 µg/mL) and repeat on 3 independent substrates | Establish LOD and fabrication reproducibility |
| Detection only at 500 µg/mL | Increase integration time to 60 s and repeat the series | Sensitivity-limited, not chemistry-limited |
| No LTA signal but R6G strong | Cysteamine-functionalised arm: bare Ag vs cysteamine-SAM Ag, LTA at 500 µg/mL on both | LTA's phosphodiester backbone is negatively charged; a positively charged -NH₃⁺ surface is a rational, LTA-specific capture strategy |
| No LTA and weak R6G | Re-fabricate substrates; do not interpret the LTA result | Substrate failure, not analyte failure |

## 7.3  EXPERIMENT 3 — Cross-analyte comparison (later)

Once LPS (Priority 1) and LTA (Priority 2) have each been tested individually, run all three analytes (LPS, LTA, NAM) on substrates from a single fabrication batch, at matched concentration and identical settings. This is the experiment that answers 'which biomarker should this platform target' with internally consistent data — but it is only meaningful once at least one analyte is known to give signal.

# 8. Raman Acquisition and Measurement Matrix

## 8.1  MANDATORY pre-analysis step — despiking

CAUTION — Before ANY peak call: reject features narrower than 3 pixels (~5 cm-1). Verify pixel width explicitly. The 14 Aug 'LTA signal' at 1095 cm-1 was a 1-pixel artifact at pixel 640 that produced a convincing SNR of 7-9. Check whether pixel 640 is a persistent hot pixel on this detector by examining a dark/blank spectrum; if so, flag that pixel permanently in the analysis pipeline.

## 8.2  Acquisition settings

| Parameter | Setting | Justification |
| --- | --- | --- |
| Wavelength | 785 nm | Unchanged; matches all prior work |
| Spectral range | 400-1800 cm⁻¹ | Covers all phosphate, lipid and saccharide targets |
| Integration | 30 s (start) | Longer than prior runs; LTA expected weak. Your raw-Si calibration showed SNR scales better than sqrt(t) on this instrument, so long integration is efficient here |
| Accumulations | 3 | Consistent with all prior work |
| Power level | 10 (highest previously used) | Check for burning at 30 s before the full run |
| Spots | 9 per concentration | Matches the published LPS reproducibility standard |
| Drop volume | 2.0 µL | Matches published LPS protocol; limits spreading on superhydrophilic black-Si |
| Drop confinement | Hydrophobic barrier ring, mandatory | Black-Si is superhydrophilic — 3 µL previously spread across an entire piece |

## 8.3  Full measurement matrix

| Order | Sample | Spots | Purpose |
| --- | --- | --- | --- |
| 1 | Blank — bare etched substrate, same session | 3 | Background reference. MANDATORY FIRST. |
| 2 | R6G 0.01 mM | 3 | SERS-activity gate — do not proceed if this fails |
| 3 | LTA 500 µg/mL | 9 | Primary detection attempt |
| 4 | LTA 100 µg/mL | 9 | Primary working concentration |
| 5 | LTA 10 µg/mL | 5 | Sensitivity |
| 6 | LTA 1 µg/mL | 5 | Sensitivity / LOD approach |
| 7 | Blank on 2nd substrate | 3 | Blank reproducibility |
| 8 | LTA 500 µg/mL on 2nd substrate | 5 | Substrate-to-substrate reproducibility |

Total ≈42 spectra. At 30 s × 3 accumulations plus repositioning, budget approximately 3.5-4 hours.

## 8.4  Detection criteria

| Verdict | Requirement |
| --- | --- |
| DETECTED | ≥1 phosphate band (1080-1095 or 1225-1250 cm⁻¹) AND ≥1 lipid band (1450 or 1131 cm⁻¹), each SNR ≥3, each ≥3 pixels wide, reproducible in ≥3 of 9 spots at consistent position (±5 cm⁻¹) |
| POSSIBLY DETECTED | Phosphate band alone at SNR ≥3 in ≥3 spots with no lipid band — record, repeat, do NOT report as detection (see the specificity limitation, Section 2.2) |
| NOT DETECTED | No band clears SNR ≥3 with ≥3-pixel width reproducibly |

# 9. File Naming and Logging

Format: EXP-LTA-{conc}-{TYPE}-S{spot}

| Example | Meaning |
| --- | --- |
| EXP-LTA-AG20-BLANK-S01 | 20 mM AgNO₃ substrate, blank, spot 1 |
| EXP-LTA-AG20-R6G-S01 | R6G control, spot 1 |
| EXP-LTA-500UGML-S01 | LTA 500 µg/mL, spot 1 |
| EXP-LTA-100UGML-S07 | LTA 100 µg/mL, spot 7 |
| EXP-LTA-1UGML-S05-SUB02 | LTA 1 µg/mL, spot 5, second substrate |

Log one row per spectrum with: Date | Experiment ID | Substrate ID | Fabrication date | Fab-to-measure interval | AgNO₃ conc/mass | HF vol | Bath vol | Etch time | Si wafer ID | LTA source/lot | LTA solution prep date | Vortex+sonicate done (Y/N) | Working conc | Drop volume | Dry method | Dry time | Power level | Integration | Accumulations | Spot # | Sample type | File name | Phosphate band (pos/height/SNR/px width) | Lipid band (pos/height/SNR/px width) | Notes

# 10. References

| Source | What it provides | Tier |
| --- | --- | --- |
| Sigma-Aldrich, Product Information Sheet, Lipoteichoic acid from Staphylococcus aureus, Product L2515, CAS 56411-57-5. | Confirmed structure (phosphodiester-linked glycerol phosphate, lipid-bound, C2 glycosylated/D-alanylated, amphipathic); solubility 5 mg/mL in water, hazy solution; storage 2-8 °C solid; aqueous stability 2 days at 4 °C or 2 months frozen | DIRECT — reagent data |
| Wu, X.; Zhao, Y.; Zughaier, S.M. (2021). Highly Sensitive Detection and Differentiation of Endotoxins Derived from Bacterial Pathogens by Surface-Enhanced Raman Scattering. Biosensors 11(7):234. DOI: 10.3390/bios11070234 | Full label-free SERS protocol on a silver substrate (2 µL drop, vacuum dry, 785 nm, 400-1800 cm⁻¹, 9 spots); peak assignment table including 850 (C-O-C saccharide), 1131 (C-C fatty acid), 1450 (CH₃/CH₂); uses LTA as a reference material | DIRECT for method; RELATED for LTA specifically |
| Literature on bacterial cell-wall Raman assignments (see Section 2 evidence tiers) | ~1080 cm⁻¹ assigned to symmetric PO₂⁻ stretch present in nucleic acids, LPS and teichoic acids; 1225-1250 cm⁻¹ antisymmetric PO₂⁻ | LITERATURE-CONFIRMED for functional group, not LTA-specific |
| Characterization and Discrimination of Gram-Positive Bacteria Using Raman Spectroscopy with PCA. Nanomaterials 7(9):248. DOI: 10.3390/nano7090248 | Gram-positive bacterial Raman characterisation context | RELATED |

NOT CONFIRMED — No publication reporting a complete, label-free SERS spectrum of PURIFIED LTA with a full peak assignment table was located, despite targeted searching. LTA appears in the SERS literature primarily as a control material. This means a well-executed LTA spectrum from your substrate would itself be a modest novel contribution — smaller than the NAM gap, but real.