Alternative Ag/Si Substrate Architectures for LPS SERS Detection

Literature-Based Fabrication Strategy — Beyond One-Step AgNO₃/HF MACE

____________________________________________________________

Sukesh P (M24IM1007)  |  IIT Jodhpur & AIIMS Jodhpur Joint IDRP  |  16 August 2026

Target: E. coli O111:B4 LPS (Sigma L2630, CAS 93572-42-0)  |  Instrument: B&W Tek i-Raman Plus BWS465-785H, 785 nm

KEY — HEADLINE FINDING — Both of the strongest published LPS SERS studies (Wu 2021; Yang 2022) used SILVER NANOROD (AgNR) ARRAYS made by oblique-angle deposition (OAD) in an electron-beam evaporator. That specific substrate is NOT replicable in your lab without e-beam evaporation. However, the REASON it works is replicable: in an AgNR array the silver morphology is engineered independently of the underlying support, producing dense, uniform, vertically-separated Ag structures with controlled nanogaps. Your one-step AgNO₃/HF process couples Si etching and Ag growth into a single uncontrolled event. The recommended route (Part G) decouples them.

CAUTION — Evidence labelling used throughout: [DIRECTLY REPORTED] = stated in the cited paper. [INFERRED FROM LITERATURE] = reasonable extrapolation across cited sources, not directly tested. [OUR PROPOSED MODIFICATION] = our design choice, no literature precedent for that exact step. No experimental evidence is invented anywhere in this document.

# A. Literature Map

## A.1  Search scope and honest limitations

Sources searched: PubMed/PMC, MDPI, RSC, ACS, ScienceDirect, Springer, OSTI, ResearchGate indexes. Searches targeted label-free LPS/endotoxin SERS, Ag/Si and Ag/black-Si SERS, SiNW SERS, porous-Si SERS, two-step MACE, photonic reduction, and MACE morphology control.

NOT CONFIRMED — LIMITATION — Several papers below were verified from abstract, PMC full text or publisher landing pages, not from a downloaded PDF of every full text. Where a numerical claim could not be traced to the source text, it is marked as such rather than asserted. Paywalled items (ACS, Elsevier, RSC full texts) were the main constraint.

## A.2  The literature divides into three groups

| Group | What it contains | Relevance to you |
| --- | --- | --- |
| Group 1 — LPS SERS benchmarks (the target) | Wu 2021; Yang 2022; DeepRaman 2025; AuNP LPS work | Defines what LPS detection looks like and what performance is achievable. Uses AgNR/AuNP, NOT Ag/Si |
| Group 2 — Ag/Si SERS fabrication (the method) | Kochylas 2021; Bai 2019; SiNW+AgNP galvanic work; porous-Si+Ag work | Defines how to build a high-EF Ag/Si substrate. Almost NONE of these tested LPS |
| Group 3 — the gap | — | No paper found combining E. coli O111:B4 LPS + Ag/black-Si + alternative MACE + 785 nm. See Part K |

KEY — The two groups barely overlap. The LPS people use evaporated AgNR; the Ag/Si people validate with R6G/dyes, not LPS. Bridging that gap IS your project's defensible contribution — but it also means you cannot simply follow one paper end-to-end.

# B. LPS SERS Benchmark Table

## B.1  Benchmark Paper 1 — Wu, Zhao & Zughaier (2021)

Wu, X.; Zhao, Y.; Zughaier, S.M. "Highly Sensitive Detection and Differentiation of Endotoxins Derived from Bacterial Pathogens by Surface-Enhanced Raman Scattering." Biosensors 2021, 11(7), 234. DOI: 10.3390/bios11070234  |  PMC8301963

| Parameter | Reported value | Tier |
| --- | --- | --- |
| Substrate type | Silver nanorod (AgNR) array on glass | [DIRECTLY REPORTED] |
| Silver morphology | Aligned Ag nanorods, 2000 nm Ag deposited at 86° incidence | [DIRECTLY REPORTED] |
| Fabrication | Oblique angle deposition (OAD) in e-beam evaporator: glass cleaned with piranha (80% H₂SO₄ / 20% H₂O₂), 20 nm Ti layer (~0.2 nm/s), 200 nm Ag film (~0.3 nm/s), then substrate tilted to 86° and Ag deposited at ~0.3 nm/s to 2000 nm; QCM monitored | [DIRECTLY REPORTED] |
| Silicon involvement | NONE — glass slides, no silicon anywhere | [DIRECTLY REPORTED] |
| LPS preparation | Stocks in pyrogen-free water at 10 nmole/mL, extensive vortexing AND sonication prior to each dilution; enteric LPS further purified, phospholipids removed by repeated 9:1 ethanol/water extraction | [DIRECTLY REPORTED] |
| LPS concentration | 10 nmole/mL (stated as corresponding to pathophysiological concentrations in clinical infection) | [DIRECTLY REPORTED] |
| Sample volume | 2 µL dispensed directly onto the substrate | [DIRECTLY REPORTED] |
| Drying | Vacuum-dried in a custom vacuum chamber (to shorten prep time) | [DIRECTLY REPORTED] |
| Raman instrument | Enwave ProRaman-L-785A2, 785 nm near-IR diode | [DIRECTLY REPORTED] |
| Laser power | 100 mW ON THE SAMPLE | [DIRECTLY REPORTED] |
| Integration time | 10 to 30 s | [DIRECTLY REPORTED] |
| Number of spectra | 9 randomly selected spots per sample, averaged | [DIRECTLY REPORTED] |
| Spectral range | 400-1800 cm⁻¹ | [DIRECTLY REPORTED] |
| Processing | Broad fluorescence background present; all baselines corrected first, then the 9 spectra averaged | [DIRECTLY REPORTED] |
| E. coli LPS included? | YES — E. coli, S. typhimurium, S. minnesota, V. cholerae, R. CE3, R. NGR, plus N. meningitidis LPS/KDO2-lipid A/lipid A | [DIRECTLY REPORTED] |
| Detection limit | NOT reported as a concentration LOD. All samples measured at a single 10 nmole/mL concentration; the paper's aim was differentiation, not LOD determination | [DIRECTLY REPORTED — absence] |
| Reproducibility | AgNR substrate spot-to-spot variation previously reported as <10%; 9 spectra shown in their Appendix C | [DIRECTLY REPORTED] |
| Differentiation method | PCA on Savitzky-Golay-derivatised, normalised, mean-centred spectra | [DIRECTLY REPORTED] |

### Can this be adapted to an Ag/Si platform?

Partly. The SAMPLE-SIDE protocol (2 µL, vortex+sonicate, dry, 785 nm, 400-1800 cm⁻¹, 9 spots, baseline-correct-then-average) transfers directly and should be copied. The SUBSTRATE-SIDE cannot be copied without an e-beam evaporator. Critically, their 100 mW at sample is roughly 3x higher than your instrument class delivers (~30 mW nominal at probe for BWS465-785x) — you must compensate with longer integration.

## B.2  Benchmark Paper 2 — Yang et al. (2022)

Yang, Y. et al. "Differentiation and classification of bacterial endotoxins based on surface enhanced Raman scattering and advanced machine learning." Nanoscale 2022, 14(24), 8806-8817. DOI: 10.1039/D2NR01277D

| Parameter | Reported value | Tier |
| --- | --- | --- |
| Substrate architecture | Silver nanorod (AgNR) array — same platform family as Wu 2021 | [DIRECTLY REPORTED] |
| Silver nanostructure | Aligned Ag nanorods | [DIRECTLY REPORTED] |
| Fabrication method | Oblique angle deposition (same group's established AgNR platform) | [DIRECTLY REPORTED] |
| LPS sample types | Eleven bacterial endotoxins | [DIRECTLY REPORTED] |
| Amount per measurement | Average detection amount 8.75 pg per measurement | [DIRECTLY REPORTED] |
| Machine learning | SVM, k-nearest neighbour, random forest, plus a modified deep-learning algorithm 'RamanNet' | [DIRECTLY REPORTED] |
| Classification performance | Most conventional ML >99% differentiation accuracy; RamanNet 100% | [DIRECTLY REPORTED] |
| Raman settings / spectral processing detail | NOT VERIFIED — full text is paywalled (RSC). Landing page and OSTI record consulted | [NOT VERIFIED] |
| Characteristic peaks | NOT VERIFIED from full text | [NOT VERIFIED] |

KEY — The 8.75 pg per measurement figure is the single most useful benchmark number in this document. It tells you the absolute MASS on the substrate that a top-tier Ag SERS platform needs — not a solution concentration. In your 2 µL deposition, 8.75 pg corresponds to roughly 4.4 ng/mL. Your planned 100 µg/mL working solution delivers ~200 ng in 2 µL — about 23,000x more material than their detection amount. You have an enormous margin, which is exactly why LPS is the right first target.

NOT CONFIRMED — A follow-up exists: 'DeepRaman: Implementing SERS together with cutting-edge machine learning for the differentiation and classification of bacterial endotoxins', Heliyon 2025, PMC11870271 — same research direction. Not analysed in detail here; relevant later if you pursue ML classification.

## B.3  Other LPS SERS evidence

| Source | Substrate | Key claim | Tier |
| --- | --- | --- | --- |
| PMC10419157 — 'A SERS-Based Biosensor for the Detection of Biological Macromolecules: The Case of the LPS Endotoxin Molecules' | 50 nm gold nanoparticles, label-free, LPS adsorbed directly | Quantitative evaluation of LPS on AuNPs 'down to few attograms'; intrinsic LPS spectral features observed with no label | [DIRECTLY REPORTED] |
| AgNP-laden black-phosphorus SERS platform (ScienceDirect S2635099823001577) | AgNPs grown in situ on black phosphorus flakes | Sepsis biomarker detection, LOD reported ~10 fM | [REPORTED — but verify the analyte identity before citing as an LPS paper] |
| PMC8530015 — 'SERS Sensing of Bacterial Endotoxin on Gold Nanoparticles' | Gold nanoparticles | Endotoxin SERS on AuNP | [DIRECTLY REPORTED] |

CAUTION — Per your Rule 3, note explicitly: the AgNP/black-phosphorus 'sepsis biomarker' paper appeared in searches for LPS but its primary analytes should be confirmed from the full text before you cite it as LPS evidence — sepsis-biomarker papers frequently target procalcitonin or interleukins rather than LPS. Do not cite it in a thesis without that check.

# C. Alternative MACE / Si Nanostructuring Methods

## C.1  Master comparison table

Abbreviated column set for readability; full per-method detail follows in C.2. All entries are Ag-compatible. Your existing process is included for comparison and clearly marked.

| # | Method | Si morphology produced | Ag morphology | Etch chemistry / duration | Extra Ag step needed? | Reported EF / LOD | Status |
| --- | --- | --- | --- | --- | --- | --- | --- |
| M0 | One-step AgNO₃/HF MACE, 2 min | Shallow black-Si / porous | In-situ dendrites + particles, uncontrolled | AgNO₃ + HF, 2 min, RT | No | Your R6G worked; LPS untested | [ALREADY TRIED BY ME] |
| M1 | Two-step MACE: Ag plate → add H₂O₂ to etch | SiNW arrays, deeper and more ordered | In-situ, then typically stripped | AgNO₃/HF plating, then HF/H₂O₂ | Usually yes | SiNW arrays demonstrated | [NEW] |
| M2 | MACE → HNO₃ Ag strip → controlled AgNO₃/HF re-deposition | SiNW/nanostructure retained | Re-grown AgNPs at controlled density | HNO₃ strip, then short AgNO₃/HF dip | Yes (that IS the step) | Density controllable | [NEW — RECOMMENDED] |
| M3 | MACE + photonic (UV/laser) reduction of Ag | SiNW array | AgNPs formed photochemically on the wires | MACE, then photoreduction | Yes | EF ≈ 1×10⁸; reusable | [NEW] |
| M4 | HF/H₂O₂ MACE with controlled oxidant ratio | NWs / pillars / porous / hierarchical, tunable by ratio | In-situ or added | e.g. 4.8 M HF + 0.1 M H₂O₂, 30 min RT | Often yes | Morphology control demonstrated | [NEW] |
| M5 | Galvanic displacement Ag coating on pre-made SiNWs | SiNWs made first | Ag shell/particles by galvanic displacement | Separate steps | Yes | EF up to 2.3×10⁸; LOD 10⁻¹⁴ M R6G reported | [NEW] |
| M6 | Porous Si + Ag immersion plating | Mesoporous Si layer | AgNPs in and on pores | PS formation, then 10 mM AgNO₃ 5 min | Yes | EF ~1.2-1.3×10⁶; NIR-SERS on Ag/porous-Si demonstrated | [NEW] |
| M7 | AgNP pre-deposition then MACE | Morphology set by initial Ag size/coverage | Pre-formed, controlled | Pre-deposit, then etch | Optional | Morphology-controlled | [NEW] |
| M8 | KOH/IPA texturing + MACE (hierarchical) | Dual-scale: pyramids + nanoscale | Ag on hierarchical surface | KOH/IPA, then MACE | Yes | Dual-scale concept | [NEW — least literature support for SERS] |
| M9 | Sintering / thermal reorganisation of Ag after MACE | SiNW retained | Ag reorganised, remanence removed | Thermal | Depends | Ag remanence removal demonstrated | [NEW — supporting step, not standalone] |

## C.2  Method detail and sourcing

| Method | Supporting source | What it actually demonstrated |
| --- | --- | --- |
| M1 Two-step MACE | Formation of silicon nanowire packed films from metallurgical-grade silicon powder using a two-step MACE method. PMC4209156 / PubMed 25349554 | Step 1: electroless Ag plating by dipping Si in aqueous AgNO₃ + HF. Step 2: H₂O₂ added to the plating solution to drive Si dissolution/etching. [DIRECTLY REPORTED] |
| M2 Ag strip + re-deposition | Reported in the SiNW/MACE literature: dendrites removed in HNO₃ aqueous solution, specimens then re-immersed in AgNO₃/HF to achieve a desired density of Ag nanostructures. Also: removal of Ag by HNO₃ + DI cleaning followed by sintering (PMC8683431) | Ag density becomes an independently controllable variable after the Si morphology is fixed. [DIRECTLY REPORTED] |
| M3 Photonic reduction | Bai, S.; Du, Y.; Wang, C.; Wu, J.; Sugioka, K. Nanomaterials 2019, 9(11), 1531. DOI: 10.3390/nano9111531 ¦ PMC6915563 | SiNW array by MACE, then photonic reduction to coat with AgNPs. EF ≈ 1×10⁸. Tested with R6G and dopamine hydrochloride. Robust enough for repeated reuse. [DIRECTLY REPORTED] |
| M4 HF/H₂O₂ ratio control | MACE process conditions reported as 4.8 M HF + 0.1 M H₂O₂, 30 min at room temperature; etching determined mainly by HF and H₂O₂ volumes | Oxidant ratio governs resulting morphology. [DIRECTLY REPORTED for the specific condition; morphology-tuning is INFERRED FROM LITERATURE] |
| M5 Galvanic displacement | 'Fabrication of silver-coated silicon nanowire arrays for SERS by galvanic displacement processes', Appl. Surf. Sci. (S0169433209012318); 'Silicon Nanowires Coated with Silver Nanostructures as Ultrasensitive Interfaces for SERS', ACS Appl. Mater. Interfaces (am900087s); 'Reusable and long-life 3D Ag nanoparticles coated Si nanowire array', Appl. Surf. Sci. (S0169433219321166) | SiNW/AgNP composite interfaces with reported LOD 10⁻¹⁴ M and EF 2.3×10⁸ for R6G; separately, 20 wt% AgNP loading on a 10 µm SiNW array gave the highest hot-spot density with EF up to 1×10⁶. [DIRECTLY REPORTED — full texts paywalled, figures verified from abstracts] |
| M6 Porous Si + Ag | 'Porous Silicon Covered with Silver Nanoparticles as SERS Substrate for Ultra-Low Concentration Detection' (PubMed 26556231); 'Porous Silicon Photonic Crystals Coated with Ag Nanoparticles', Nanomaterials 2018, 8, 872, DOI 10.3390/nano8110872; 'Near-Infrared SERS on Silver-Coated Porous Silicon Photonic Crystals' PMC6473976 | Maximum R6G SERS enhancement reported for immersion plating from 10 mM aqueous AgNO₃ for 5 minutes. EF ~1.2×10⁶ at 10⁻⁶ mol/L; crystal violet detectable at 100 pM with EF 1.3×10⁶. NIR (785-relevant) SERS on Ag/porous-Si demonstrated. [DIRECTLY REPORTED] |
| M0 (yours) reference point | Kochylas, I.; Gardelis, S.; Likodimos, V.; Giannakopoulos, K.P.; Falaras, P.; Nassiopoulou, A.G. Nanomaterials 2021, 11(7), 1760. DOI: 10.3390/nano11071760 | One-step MACE in AgNO₃/HF giving SiNWs decorated with Ag aggregates or dendrites; EF near 1×10⁸ and R6G LOD to 10⁻¹³ M reported. Confirms your architecture family is capable — but this is the process you have already run. [DIRECTLY REPORTED] |

KEY — Note on M0/Kochylas: one-step AgNO₃/HF CAN reach EF ~10⁸ in expert hands. Your R6G result already showed strong enhancement. So your problem is probably NOT that one-step MACE is incapable — it is that the Ag morphology it produces is uncontrolled and unoptimised for a weak, non-resonant analyte. That is the specific deficiency the recommended method addresses.

# D. Top 5 Candidate Fabrication Approaches for YOUR Lab

Ranked against your stated constraints: Si available, Ag chemistry available, HF available under institutional procedure, 785 nm Raman, R6G available, no expensive lithography, no months of optimisation, SEM/EDS NOT assumed.

## #1 BEST — M2: Two-step MACE with Ag strip and controlled re-deposition

| Aspect | Detail |
| --- | --- |
| Why this rank | Decouples Si morphology from Ag morphology — the single most important structural difference from your current process, and the property that makes the AgNR benchmark substrates work. Uses ONLY chemicals you already have (AgNO₃, HF, plus HNO₃). No new equipment. Ag density becomes an independent, tunable variable. Fabrication adds roughly 30-60 minutes to your existing workflow. |
| Main risk / cost | HNO₃ is the only new reagent. Requires care not to over-strip (which would also attack the Si nanostructure). |
| Realistic fabrication time | 1 day |

## #2 — M5: Pre-made SiNWs + galvanic-displacement Ag coating

| Aspect | Detail |
| --- | --- |
| Why this rank | Strongest reported numbers of any Ag/Si route found (EF 2.3×10⁸, R6G LOD 10⁻¹⁴ M), and mechanistically similar to #1 but with the Ag step performed as a deliberate, separately-optimised deposition rather than a re-deposition. Chemically identical toolkit to what you already use. |
| Main risk / cost | Full texts paywalled — exact deposition parameters not verified. You would be optimising the Ag step semi-blind. |
| Realistic fabrication time | 1-2 days |

## #3 — M6: Porous Si + Ag immersion plating

| Aspect | Detail |
| --- | --- |
| Why this rank | Has a CONCRETE, DIRECTLY-REPORTED optimal recipe you can copy verbatim: immersion plating from 10 mM aqueous AgNO₃ for 5 minutes gave maximum R6G enhancement. Porous Si also offers a high-surface-area adsorption environment, plausibly favourable for an amphiphilic macromolecule like LPS. Near-IR SERS on Ag/porous-Si is specifically demonstrated, which matters for your 785 nm instrument. |
| Main risk / cost | Requires forming porous Si first (electrochemical anodisation in HF, or stain etching). Anodisation needs a cell and current source; stain etching is simpler but less controlled. |
| Realistic fabrication time | 1-2 days |

## #4 — M4: HF/H₂O₂ MACE with controlled oxidant ratio

| Aspect | Detail |
| --- | --- |
| Why this rank | Cleanest single-variable DOE: keep Ag deposition fixed, vary the H₂O₂:HF ratio to sweep through nanowire / pillar / porous / hierarchical morphologies. A reported working condition exists (4.8 M HF + 0.1 M H₂O₂, 30 min, RT) — and note your existing bath is already ~4.8 M HF, so only the oxidant changes. |
| Main risk / cost | H₂O₂ is a new reagent in the etch bath. 30 min etch is much longer than your 2 min — a genuinely different process, which is the point. |
| Realistic fabrication time | 1 day |

## #5 — M3: MACE + photonic reduction

| Aspect | Detail |
| --- | --- |
| Why this rank | Highest verified EF of the alternatives (≈1×10⁸) and explicitly reusable, from a well-documented open-access paper (Bai 2019, DOI 10.3390/nano9111531). Conceptually elegant: Si structure first, Ag formed photochemically afterwards. |
| Main risk / cost | Requires a controlled UV/laser photoreduction setup. The exact optical parameters were not verified from the full text in this pass. Highest equipment uncertainty of the five. |
| Realistic fabrication time | 2-3 days + setup |

# E. Two Proposed New Ag/Si Substrate Designs

## E.1  NEW SENSOR A — 'Decoupled Two-Step Ag/SiNW' (primary)

| Field | Specification |
| --- | --- |
| Principle | Separate the two jobs that your current process does simultaneously. Step 1 builds the silicon nanostructure. Step 2 removes all of that Ag. Step 3 grows fresh Ag under short, controlled, independently-variable conditions. Ag size/density then becomes a knob you can turn without changing the Si. |
| Fabrication sequence | (1) Clean Si: acetone → IPA → DI, 5 min sonication each. (2) Native-oxide strip: 4.1 mL 49% HF + 15.9 mL DI (20 mL, ~10% HF), 30-60 s. (3) MACE etch: 0.170 g AgNO₃ + 8.6 mL HF, to 50 mL, 10 MINUTES (deliberately longer than your 2 min, to build a deeper nanostructure). (4) Rinse. (5) Ag STRIP: immerse in dilute HNO₃ until the surface visibly loses its dark metallic Ag deposit; rinse thoroughly in DI. (6) Ag RE-DEPOSITION: short dip in dilute AgNO₃/HF — start at 5 mM AgNO₃ with 1 mL 49% HF in 50 mL, 60 s. (7) Rinse, N₂ dry. |
| Why good hotspots | In the re-deposition step, Ag⁺ nucleates on a Si surface that already has high aspect-ratio topography. Nucleation is brief and dilute, favouring many small discrete particles rather than large coalesced dendrites — the regime associated with dense sub-5 nm interparticle gaps. [INFERRED FROM LITERATURE] |
| Expected Si morphology | Deeper nanowire/porous nanostructure than your 2-min result, from the 10-min etch. [INFERRED FROM LITERATURE] |
| Expected Ag morphology | Small, discrete AgNPs distributed over the nanostructure walls, at a density set by the re-deposition time/concentration. [INFERRED FROM LITERATURE] |
| Why favourable for LPS | LPS is amphiphilic: lipid A drives physisorption onto Ag, and the O-antigen chain is hydrophilic. A surface with many small, accessible Ag particles on an open nanostructure gives more adsorption sites within field-enhancement range than a surface dominated by a few large dendrites. [INFERRED FROM LITERATURE] |
| Literature support | Ag strip in HNO₃ and re-immersion in AgNO₃/HF to set Ag density is [DIRECTLY REPORTED] in the SiNW/MACE literature. The specific 10-min etch / 5 mM / 60 s parameters are [OUR PROPOSED MODIFICATION]. |
| Expected advantages | Uses only AgNO₃, HF, HNO₃. No new equipment. Ag becomes an independent variable. Directly comparable to your existing substrate since the Si chemistry family is unchanged. |
| Expected failure modes | (a) HNO₃ over-strips and damages/oxidises the Si nanostructure. (b) Re-deposition too short → too little Ag → weak signal. (c) Re-deposition too long → recreates the dendritic morphology you already have. Mitigate by running 3 re-deposition times in parallel (30 s / 60 s / 120 s). |
| Fabrication time | Approximately 2-2.5 hours for a batch of 3-4 substrates. |
| Minimum characterisation | REQUIRED: visual inspection (surface should be dark/black after etch; visibly lighter after HNO₃ strip; slightly darkened again after re-deposition) + R6G SERS as the functional test. PREFERRED: SEM. OPTIONAL: EDS, UV-Vis, AFM, XPS. |

## E.2  NEW SENSOR B — 'HF/H₂O₂ Oxidant-Controlled MACE' (backup / parallel arm)

| Field | Specification |
| --- | --- |
| Principle | Change the OXIDANT rather than the metal. In your current process Ag⁺ is both catalyst and oxidant, so Ag consumption and Si etching are inseparable. Adding H₂O₂ as a separate oxidant lets etching continue after Ag deposition is complete, giving deeper, more ordered nanostructures at a controllable rate. |
| Fabrication sequence | (1) Clean and oxide-strip as above. (2) Ag seeding: brief dip in dilute AgNO₃/HF (5 mM AgNO₃ + 1 mL HF in 50 mL, 60 s) to deposit catalyst particles only. (3) Rinse. (4) Etch in HF/H₂O₂: target ~4.8 M HF + 0.1 M H₂O₂ (a reported working condition), room temperature. Run a time series: 5 / 15 / 30 min. (5) Rinse, N₂ dry. (6) Optional Ag top-up dip if R6G signal is weak. |
| Why good hotspots | Decoupling catalyst deposition from etch progression yields more vertical, ordered nanowire arrays; Ag particles sit at wire tips and in inter-wire gaps. [INFERRED FROM LITERATURE] |
| Expected Si morphology | Ordered SiNW array, depth increasing with etch time. [INFERRED FROM LITERATURE] |
| Expected Ag morphology | Discrete catalyst particles retained from the seeding step, positioned at etch fronts. [INFERRED FROM LITERATURE] |
| Why favourable for LPS | Deep, open inter-wire channels give a large accessible surface for a large amphiphilic molecule; 785 nm penetrates several hundred nm into Si, so buried structure is optically reachable. [INFERRED FROM LITERATURE] |
| Literature support | 4.8 M HF + 0.1 M H₂O₂, 30 min, RT is [DIRECTLY REPORTED]. The seeding step and time series are [OUR PROPOSED MODIFICATION]. |
| Expected advantages | Cleanest single-variable design; only ONE new reagent (H₂O₂); HF molarity unchanged from your current bath. |
| Expected failure modes | (a) Too little Ag retained → no SERS despite good Si structure (fix with the top-up dip). (b) 30 min etch may over-etch thin pieces. (c) H₂O₂ + HF + Ag is a more vigorous system — see safety note. |
| Fabrication time | Approximately 1.5 hours plus etch time (up to 30 min per piece). |
| Minimum characterisation | REQUIRED: visual + R6G SERS. PREFERRED: SEM. OPTIONAL: everything else. |

CAUTION — SAFETY — Sensor B adds H₂O₂ to an HF system. Do NOT pre-mix concentrated H₂O₂ and HF and store it. Prepare the etch bath immediately before use, in PP/PTFE, in a fume hood, with all standard HF precautions (double gloves with HF-rated outer layer, face shield, calcium gluconate gel in date, never work alone). Sensor A's HNO₃ strip must also be done in the fume hood, and HNO₃ must never be mixed with organics.

# F. Comparison Against Your Existing 2-Minute Process

| Attribute | M0 — Your current process [ALREADY TRIED] | Sensor A [NEW] | Sensor B [NEW] |
| --- | --- | --- | --- |
| Etch chemistry | AgNO₃ + HF, Ag⁺ is the only oxidant | AgNO₃ + HF, then HNO₃, then fresh AgNO₃/HF | AgNO₃ seed, then HF + H₂O₂ |
| Etch duration | 2 min | 10 min etch + 1 min re-deposition | 5-30 min |
| Si and Ag control | COUPLED — one event sets both | DECOUPLED — three separate steps | DECOUPLED — seeding separate from etching |
| Ag morphology control | None (whatever forms, forms) | Direct: via re-deposition time and concentration | Indirect: via seeding step |
| New reagents needed | — | HNO₃ | H₂O₂ |
| New equipment needed | — | None | None |
| Genuinely different from M0? | — | YES — the strip/re-deposit sequence has no analogue in your process | YES — different oxidant, different duration, seeded catalyst |
| Added time vs. current | — | ≈ +1 hour | ≈ +30 min |

# G. Recommended Route — Single Answer

KEY — RECOMMENDATION: Build NEW SENSOR A — two-step MACE with HNO₃ Ag strip and controlled AgNO₃/HF re-deposition. Run three re-deposition times (30 s / 60 s / 120 s) in a single session, validate all three with R6G, then take the best one straight to LPS.

| Question | Answer |
| --- | --- |
| 1. Which Si fabrication method? | Two-step MACE: 10-min AgNO₃/HF etch → HNO₃ Ag strip → short controlled AgNO₃/HF Ag re-deposition (Sensor A, Section E.1). |
| 2. Why better than the old 2-min process? | Your current process couples Si etching and Ag growth into one uncontrolled event, so Ag morphology is whatever the etch happens to leave. Both top LPS benchmark papers succeeded on substrates where Ag morphology was engineered independently of the support. Sensor A reproduces that decoupling using only chemistry you already have. |
| 3. Expected morphology? | Deeper SiNW/porous Si from the longer etch, decorated with small discrete AgNPs at a density set by re-deposition time. [INFERRED FROM LITERATURE] |
| 4. How are Ag hotspots generated? | By brief, dilute Ag⁺ nucleation onto pre-existing high-aspect-ratio Si topography — favouring many small discrete particles with narrow interparticle gaps rather than coalesced dendrites. |
| 5. Why good for LPS? | LPS is amphiphilic; lipid A physisorbs to Ag while the polysaccharide chain stays hydrated. Many small accessible Ag particles on an open nanostructure maximise the number of LPS molecules sitting within the steep field-enhancement zone. |
| 6. Strongest LPS benchmark? | Wu, Zhao & Zughaier 2021 (DOI 10.3390/bios11070234) for the sample-side protocol and peak assignments — it includes E. coli LPS and gives a complete, copyable measurement method. Yang 2022 (DOI 10.1039/D2NR01277D) for the performance target: 8.75 pg per measurement. |
| 7. Which Raman bands? | Primary: 989 and 1330 cm⁻¹. Supporting: 550, 680, 850, 1131, 1450 cm⁻¹. See Part I. |
| 8. Fastest validation experiment? | R6G on all three re-deposition variants, same session, same settings — a 1-hour go/no-go before any LPS is consumed. |
| 9. Mandatory controls? | Same-session bare-substrate blank; R6G positive control; LPS on unetched bare Si (negative control); solvent-only (water) drop control. |
| 10. What makes us proceed? | Any re-deposition variant giving R6G signal at least comparable to your existing 2-min substrate, PLUS LPS showing 989 and 1330 cm⁻¹ at SNR ≥3 in ≥3 of 9 spots. |
| 11. What makes us abandon this architecture? | All three variants give R6G markedly WEAKER than your existing 2-min substrate — that would mean the HNO₃ strip is damaging the structure, and Sensor B becomes the route. |
| 12. Next optimisation if it fails? | Move to Sensor B (HF/H₂O₂ oxidant control). If that also fails, go to M6 (porous Si + 10 mM AgNO₃ 5 min immersion plating), which has the most concrete copyable recipe. |

# H. LPS Detection Experiment

## H.1  Sample preparation (adapted from Wu 2021 — sample side copied deliberately)

| Step | Action | Exact quantity |
| --- | --- | --- |
| H1 | Weigh LPS (fume hood — L2630 is H300, fatal if swallowed) | 1.0 mg, recorded to 0.1 mg |
| H2 | Dissolve in pyrogen-free / high-purity water | 1.00 mL → 1 mg/mL stock |
| H3 | Vortex | 2 min |
| H4 | Sonicate | 15 min (Wu 2021 explicitly sonicated before each dilution) |
| H5 | Working series (vortex 30 s + sonicate 5 min at EACH step) | 100 / 10 / 1 / 0.1 µg/mL, 1.00 mL each |
| H6 | Deposit | 2.0 µL per spot (matches Wu 2021 exactly) |
| H7 | Dry | Vacuum desiccator preferred; air-dry acceptable if logged |

NOT CONFIRMED — Concentration is expressed in mass units only. L2630 is a heterogeneous phenol-extracted LPS with no single molecular weight — no molarity is computed anywhere in this document.

## H.2  Measurement matrix

| # | Sample | Substrate | Spots | Purpose |
| --- | --- | --- | --- | --- |
| 1 | Blank — bare Ag/Si, no drop | Sensor A (best variant) | 3 | Same-session background. MANDATORY FIRST. |
| 2 | Water-only drop, dried | Sensor A | 3 | Solvent/residue control |
| 3 | R6G 0.01 mM | Sensor A | 3 | SERS-activity gate |
| 4 | LPS 100 µg/mL | UNETCHED bare Si | 3 | Negative control — proves any signal is SERS, not ordinary Raman |
| 5 | LPS 100 µg/mL | Sensor A | 9 | Primary detection attempt |
| 6 | LPS 10 µg/mL | Sensor A | 9 | Concentration response |
| 7 | LPS 1 µg/mL | Sensor A | 5 | Concentration response |
| 8 | LPS 0.1 µg/mL | Sensor A | 5 | LOD approach |
| 9 | Blank + LPS 100 µg/mL | Sensor A, 2nd substrate | 3+5 | Substrate-to-substrate reproducibility |

Approximately 48 spectra. At 30 s × 3 accumulations, budget ~4 hours.

## H.3  Raman settings

| Parameter | Setting | Basis |
| --- | --- | --- |
| Wavelength | 785 nm | Matches Wu 2021 and your instrument |
| Range | 400-1800 cm⁻¹ | Matches Wu 2021 exactly |
| Integration | 30 s (start) | Wu used 10-30 s at 100 mW; you have less power, so use their upper bound |
| Accumulations | 3 | Your established practice |
| Power level | 10 (your highest previously used) | Check for burning before the full run |
| Spots | 9 per primary concentration | Matches Wu 2021 exactly |
| Processing | Despike → baseline correct → average the 9 | Wu baseline-corrected then averaged; despiking added per your own artifact finding |

# I. LPS Raman Fingerprint Table

| Raman shift (cm⁻¹) | Assignment | Source | LPS type | Confidence | Relevance to E. coli O111:B4 |
| --- | --- | --- | --- | --- | --- |
| 550 | β(CH₂) in ring | Wu 2021 Table 1 | N. meningitidis LPS | Measured SERS on Ag | Moderate — enteric LPS showed related bands at 543/552 |
| 680 | δ(C-O-C); fatty acid | Wu 2021 Table 1 | N. meningitidis LPS | Measured SERS on Ag | Moderate — lipid A is common to all LPS |
| 850 | ν(C-O-C); saccharides, 1,4-glycosidic link | Wu 2021 Table 1 | N. meningitidis LPS | Measured SERS on Ag | Moderate-high — O111:B4 has a long O-antigen, rich in glycosidic links |
| 988-989 | β(CH); carbohydrates | Wu 2021 Table 1 + text | N. meningitidis LPS and KDO2-lipid A | Measured SERS on Ag; explicitly called a significant peak | HIGH — enteric LPS showed a matching band at 981; PRIMARY TARGET |
| 1131 | ν(C-C); fatty acid | Wu 2021 Table 1 | N. meningitidis LPS | Measured SERS on Ag | Moderate-high — lipid A acyl chains common to all LPS |
| 1330 | δ(CH); phospholipid | Wu 2021 Table 1 + text | N. meningitidis LPS; also seen in R. CE3/R. NGR enteric spectra | Measured SERS on Ag; explicitly called a significant peak | HIGH — PRIMARY TARGET |
| 1450 | α/β(CH₃/CH₂) | Wu 2021 Table 1 | N. meningitidis LPS | Measured SERS on Ag | Moderate-high — generic lipid deformation |
| 543, 552, 981 | saccharide and lipid bands | Wu 2021 §3.2 | Enteric LPS incl. E. coli; Salmonella; Rhizobium | Measured SERS on Ag | HIGH — these are the ENTERIC values, closest to your O111:B4 |
| ~2900 (C-H stretch) | ν(C-H) lipid chains | NOT in Wu 2021 — their range stopped at 1800 | — | NOT VERIFIED for LPS SERS | CAUTION — see note below |

CAUTION — About the ~2900 cm-1 CH-stretch region: Wu 2021 collected only 400-1800 cm-1, so it provides NO evidence for LPS bands at 2900. Your own instrument's usable range extends to about 2840 cm-1 (from your raw-Si files), which only just reaches the bottom edge of the CH-stretch envelope. Do NOT plan to rely on 2900 cm-1. It is also a highly non-specific region — almost any organic contamination, fingerprint residue or pump oil gives C-H stretch signal.

KEY — Organism-specific differences are real and documented: Wu 2021 showed E. coli and V. cholerae spectra were 'entirely different' from Salmonella and Rhizobium LPS spectra, which is exactly what enabled bacterial differentiation by PCA. Therefore expect your O111:B4 spectrum to resemble the ENTERIC values (543/552/981) more closely than the N. meningitidis table values.

# J. Detection and LOD Strategy

## J.1  Objective detection criteria

| Criterion | Threshold |
| --- | --- |
| SNR | ≥3 above local flanking baseline, noise estimated from successive differences in a peak-free region within ±100 cm⁻¹ |
| Peak width | ≥3 pixels (~5 cm⁻¹). MANDATORY artifact rejection — a 1-pixel feature is a detector spike, not a band |
| Position tolerance | ±5 cm⁻¹ of target (your instrument shows a systematic ~2-3 cm⁻¹ offset: Si TO phonon found at 517.8 vs textbook 520-521) |
| Minimum spots | ≥3 of 9 spots must independently satisfy the above |
| Band combination | BOTH primary bands (989 and 1330 cm⁻¹) required — one alone is insufficient |
| Blank comparison | Neither band may be present above SNR 3 in the same-session blank or in the water-only control |
| Bare-Si control | LPS on unetched Si must NOT show the bands — otherwise you are seeing ordinary Raman or contamination, not SERS |
| Concentration response | Peak height must decrease monotonically (within error) from 100 → 10 → 1 µg/mL. A signal that does not track concentration is not the analyte |

## J.2  LOD / LOQ

Determine LOD only AFTER detection is established and only from the concentration series: LOD = the lowest concentration at which the detection criteria above are still met in ≥3 of 9 spots. Report it as a mass concentration (µg/mL) AND as mass-on-substrate (concentration × 2 µL), so it can be compared directly with the Yang 2022 benchmark of 8.75 pg per measurement.

NOT CONFIRMED — Do not compute a statistical 3σ/slope LOD from fewer than 5 concentration points with replicates. With the 4-point series specified here, report an OBSERVED lowest-detected concentration and label it as such. LOQ is not justified at this stage.

# K. Novelty Assessment

## K.1  What has already been done

LPS SERS on evaporated Ag nanorod arrays, including E. coli LPS, with full peak assignments — Wu 2021.

LPS/endotoxin SERS with machine-learning classification on AgNR, 8.75 pg per measurement — Yang 2022.

Label-free LPS SERS on gold nanoparticles down to attogram quantities — PMC10419157.

High-EF Ag/SiNW SERS substrates by one-step MACE, two-step MACE, galvanic displacement, photonic reduction and porous-Si routes — validated with R6G and other dyes.

## K.2  What was NOT found

NOT CONFIRMED — NOT FOUND IN THE SEARCHED LITERATURE — a study combining E. coli O111:B4 LPS + Ag/black-silicon (or Ag/SiNW) + an alternative two-step MACE architecture + 785 nm SERS. Stated deliberately as 'not found in the searched literature' rather than 'first ever', per Rule 3. Absence of evidence here reflects the limits of this search (paywalled full texts, non-indexed theses), not proof of absence.

## K.3  Three defensible novelty directions

| Direction | Claim you could defend | Strength |
| --- | --- | --- |
| 1. Analyte-platform bridging | First reported label-free LPS SERS detection on a MACE-fabricated Ag/black-Si substrate (as opposed to evaporated AgNR or colloidal Au/Ag) | STRONGEST — the two literatures genuinely do not overlap |
| 2. Fabrication-morphology-analyte link | Systematic comparison of coupled (one-step) vs decoupled (two-step, strip + re-deposit) Ag morphology on the SAME Si platform, evaluated with a weak non-resonant biological analyte rather than R6G | STRONG — most Ag/Si papers optimise on dye only, which your own NAM/LPS experience shows is misleading |
| 3. Low-cost platform equivalence | Demonstrating that chemically-fabricated Ag/Si approaches the performance of e-beam-evaporated AgNR for endotoxin detection | MODERATE — depends entirely on how good your numbers turn out |

# L. Materials and Equipment

## L.1  Chemicals

| Chemical | Spec | Needed for | Have it? |
| --- | --- | --- | --- |
| LPS E. coli O111:B4 | Sigma L2630, CAS 93572-42-0, 10 mg | Target analyte | Confirmed product |
| AgNO₃ | ACS ≥99% | MACE etch + Ag re-deposition | Yes |
| HF | 48-51% | Oxide strip, MACE etch | Yes (institutional procedure) |
| HNO₃ | 65-70% conc., diluted for use | Ag STRIP — Sensor A. NEW REAGENT | Need to confirm |
| H₂O₂ | 30% w/w | Oxidant — Sensor B only. NEW REAGENT | Need to confirm |
| R6G | 0.01 mM working | Positive control | Yes |
| Acetone, IPA | ACS | Cleaning | Yes |
| DI water | Type I; pyrogen-free preferred for LPS | Solvent, rinsing | Yes / confirm pyrogen-free |
| Calcium gluconate gel 2.5% | In date | HF first aid — MANDATORY | Must confirm |

## L.2  Equipment and characterisation status

| Item | Status | Note |
| --- | --- | --- |
| 785 nm Raman (BWS465-785H) | REQUIRED — have it | Core instrument |
| Fume hood, PP/PTFE labware, HF PPE | REQUIRED | Non-negotiable for all HF steps |
| Analytical balance (0.1 mg) | REQUIRED | LPS and AgNO₃ weighing |
| P2/P10 micropipette | REQUIRED | Accurate 2 µL delivery |
| Vortex mixer + bath sonicator | REQUIRED | LPS micelle dispersion — not optional |
| Hydrophobic barrier (wax/PAP) pen | REQUIRED | Black-Si is superhydrophilic; drops spread without it |
| Vacuum desiccator | PREFERRED | Wu 2021 vacuum-dried; air-dry acceptable if logged |
| SEM | PREFERRED — NOT assumed | Would confirm morphology; the R6G functional test substitutes for it |
| EDS | OPTIONAL | Would confirm Ag presence/removal after HNO₃ strip |
| UV-Vis | OPTIONAL | Would characterise plasmon resonance vs 785 nm |
| AFM / XPS | OPTIONAL | Not needed for this decision |

KEY — You can run this entire experiment with NO electron microscopy. R6G SERS intensity is your functional proxy for 'did the Ag morphology change usefully' — and since you already have R6G data from your 2-min substrate, you have a built-in comparison baseline.

# M. References

| # | Reference | DOI / link | Used for |
| --- | --- | --- | --- |
| 1 | Wu, X.; Zhao, Y.; Zughaier, S.M. Highly Sensitive Detection and Differentiation of Endotoxins Derived from Bacterial Pathogens by Surface-Enhanced Raman Scattering. Biosensors 2021, 11(7), 234. | 10.3390/bios11070234 — https://doi.org/10.3390/bios11070234 (PMC8301963) | PRIMARY LPS benchmark: full method, peak table, E. coli LPS included |
| 2 | Yang, Y. et al. Differentiation and classification of bacterial endotoxins based on surface enhanced Raman scattering and advanced machine learning. Nanoscale 2022, 14(24), 8806-8817. | 10.1039/D2NR01277D — https://doi.org/10.1039/D2NR01277D | Performance benchmark: 8.75 pg/measurement, ML classification |
| 3 | Bai, S.; Du, Y.; Wang, C.; Wu, J.; Sugioka, K. Reusable SERS Substrates Made of Silicon Nanowire Array Coated with Silver Nanoparticles Fabricated by MACE and Photonic Reduction. Nanomaterials 2019, 9(11), 1531. | 10.3390/nano9111531 — https://doi.org/10.3390/nano9111531 (PMC6915563) | Method M3: MACE + photonic reduction, EF ~1×10⁸, reusable |
| 4 | Kochylas, I.; Gardelis, S.; Likodimos, V.; Giannakopoulos, K.P.; Falaras, P.; Nassiopoulou, A.G. Improved SERS Sensitivity Using Si Nanowires/Silver Nanostructures by a Single Step MACE. Nanomaterials 2021, 11(7), 1760. | 10.3390/nano11071760 — https://doi.org/10.3390/nano11071760 (PMC8308179) | Reference point for one-step MACE capability (your architecture family) |
| 5 | Formation of silicon nanowire packed films from metallurgical-grade silicon powder using a two-step MACE method. | PMC4209156 / PubMed 25349554 | Method M1: two-step MACE (Ag plate, then add H₂O₂) |
| 6 | Removal of Ag remanence and improvement in structural attributes of silicon nanowires array via sintering. | PMC8683431 | Method M9 / Ag strip context: HNO₃ cleaning and sintering to remove Ag |
| 7 | Silicon Nanowires Coated with Silver Nanostructures as Ultrasensitive Interfaces for SERS. ACS Appl. Mater. Interfaces. | https://pubs.acs.org/doi/abs/10.1021/am900087s | Method M5: EF 2.3×10⁸, R6G LOD 10⁻¹⁴ M (abstract-level verification) |
| 8 | Fabrication of silver-coated silicon nanowire arrays for SERS by galvanic displacement processes. Appl. Surf. Sci. | https://www.sciencedirect.com/science/article/abs/pii/S0169433209012318 | Method M5: galvanic displacement Ag on SiNW |
| 9 | Reusable and long-life 3D Ag nanoparticles coated Si nanowire array as sensitive SERS substrate. Appl. Surf. Sci. | https://www.sciencedirect.com/science/article/abs/pii/S0169433219321166 | Method M5 variant: 20 wt% AgNP loading, EF up to 1×10⁶ |
| 10 | Porous Silicon Covered with Silver Nanoparticles as SERS Substrate for Ultra-Low Concentration Detection. | PubMed 26556231 | Method M6: optimal immersion plating 10 mM AgNO₃, 5 min |
| 11 | Porous Silicon Photonic Crystals Coated with Ag Nanoparticles as Efficient Substrates for Detecting Trace Explosives Using SERS. Nanomaterials 2018, 8, 872. | 10.3390/nano8110872 — https://doi.org/10.3390/nano8110872 (PMC6266845) | Method M6: EF ~1.2-1.3×10⁶ |
| 12 | Near-Infrared SERS on Silver-Coated Porous Silicon Photonic Crystals. | PMC6473976 | Method M6 at NIR — directly relevant to 785 nm |
| 13 | A SERS-Based Biosensor for the Detection of Biological Macromolecules: The Case of the LPS Endotoxin Molecules. | PMC10419157 | Label-free LPS on 50 nm AuNPs, attogram-level |
| 14 | DeepRaman: SERS with machine learning for differentiation and classification of bacterial endotoxins. Heliyon 2025. | PMC11870271 — https://www.sciencedirect.com/science/article/pii/S2405844025009302 | Follow-up ML endotoxin work |
| 15 | Sigma-Aldrich product page, Lipopolysaccharides from E. coli O111:B4, L2630. | CAS 93572-42-0; MDL MFCD00164401 | Reagent identity, solubility, hazard (H300) |

# N. DO THIS FIRST

KEY — The shortest scientifically defensible route from silicon wafer to LPS detection. Two lab days.

## DAY 1 — Fabrication (~3 hours)

| Step | Action | Quantity / time |
| --- | --- | --- |
| 1 | Cut 4 Si pieces from one wafer; record wafer ID | ~1×1 cm each |
| 2 | Clean: acetone → IPA → DI | 5 min sonication each |
| 3 | Oxide strip | 4.1 mL 49% HF + 15.9 mL DI (20 mL); 30-60 s |
| 4 | MACE etch — ALL 4 pieces together | 0.170 g AgNO₃ + 8.6 mL HF, to 50 mL; 10 MINUTES |
| 5 | Rinse thoroughly, N₂ dry | ≥30 s DI |
| 6 | SET ASIDE PIECE 1 — do NOT strip it | This is your 10-min one-step comparison control |
| 7 | HNO₃ Ag strip — pieces 2, 3, 4 | Dilute HNO₃ until dark Ag deposit visibly lifts; rinse thoroughly |
| 8 | Ag re-deposition — separate baths, different times | 5 mM AgNO₃ + 1 mL 49% HF in 50 mL. Piece 2: 30 s. Piece 3: 60 s. Piece 4: 120 s |
| 9 | Rinse, N₂ dry, label each piece immediately | — |

## DAY 2 — Validation and detection (~4 hours)

| Step | Action | Detail |
| --- | --- | --- |
| 10 | Same-session BLANK on every piece | 3 spots each, 30 s, 3 acc., power 10, 400-1800 cm⁻¹ |
| 11 | R6G 0.01 mM on all 4 pieces | 3 spots each, identical settings. THIS IS THE GO/NO-GO GATE |
| 12 | Compare R6G 1509 cm⁻¹ height across pieces 1-4 | Pick the winner. If all re-deposited pieces lose badly to piece 1, the HNO₃ strip is damaging the structure — switch to Sensor B |
| 13 | Apply hydrophobic barrier rings to the winning piece | 3-5 rings, dry 5 min |
| 14 | LPS 100 µg/mL, 2.0 µL per spot, dry | 9 spots |
| 15 | LPS negative control on unetched bare Si | 3 spots — mandatory |
| 16 | If detected: run 10, 1, 0.1 µg/mL | 5 spots each |
| 17 | Analyse: despike → baseline → per-spot criteria (Part J) → then average | Do NOT average before per-spot evaluation |

RESULT — Why this is the fastest defensible route: one fabrication session produces the control AND three variants of a genuinely new architecture; R6G gates the whole thing in under an hour before any LPS is consumed; and the sample-side protocol is copied directly from the strongest published LPS benchmark, so your results are comparable to it.